## MAIN COVER
#VALUE!

## Index
Auto+Hide+Values
Index
SMART 3G-D2
SMART 3G-C2-LD
SMART 3G-C2
SMART 3G-D3
SMART 3G-C3
SMART 3G-GRI
SMART S-SS
SMART S-MS
SMART S-MS MED
SMART S-IS
SMART 3 R
SMART 3 R BLIZZARD
SMART 3 NC
SMART 3H Lite
SMART 3H-FM Lite
PL4+ & PL4+D
MULTISCAN 8+
MULTISCAN++ S1
MULTISCAN++ S2
SECURNET++ & SCADA LITE
MULTISCAN++ MED
SMART P
MULTISCAN++ PK
Strumenti di calibrazione
Ricambi
Bombole
T&C

## Technology
Auto+Hide+Values
SENSING TECHNOLOGIES AND PERIODICAL MAINTENANCE
Catalytic sensors (Pellistors) offer excellent output linearity up to 100% LFL and have an expected lifetime of about 5 years. Their performance can be affected by the presence of certain substances that may significantly alter the sensor’s response and even cause irreversible damage. The presence of inhibitors or poisons is the most common cause of problems in gas detection; for this reason, it is essential to avoid any contamination. Among the most common poisons or inhibitors are silicones, halogens, halogenated hydrocarbons, tetraethyl lead, hydrogen sulphide, and sulphur dioxide. These compounds do not affect Infrared sensors, which are the ideal solution whenever flammable gases must be detected in environments where catalytic sensors cannot be used. This technology, based on an optical system, offers clear advantages, being independent of environmental factors. Infrared sensors have an expected lifetime of more than 5 years. Electrochemical sensors are used to detect toxic gases. Their average lifetime is about 3 years, except for sensors designed for Ammonia detection, which typically last less (1–2 years). The actual lifetime depends on the gas concentration in the environment: the higher the concentration, the shorter the lifetime. Periodic functional checks must be carried out to verify the correct operation of the detectors. In EU countries, gas functional tests and calibration procedures are required by current regulations and defined by EN 60079-29-2. This standard provides guidance on the selection, installation, use, and maintenance of gas detection systems intended for industrial and civil applications. All SENSITRON gas detectors are factory-calibrated before shipment. Alarm thresholds and measuring ranges can be adjusted in our laboratories upon customer request. The warranty does not apply to instruments that have been tampered with, repaired by unauthorized laboratories, damaged, or misused. Even during the warranty period, verification and calibration operations are excluded, as they are not covered by “free maintenance.”

## COVER SMART 3G
#VALUE!

## SMART 3G-D2
Auto+Hide+Values
SMART 3G-D2 With display for zone 1 category 2
The SMART 3G-D2 is designed to meet rigorous industrial standards, enabling the monitoring of toxic, flammable, and refrigerant gases in harsh environments and classified areas. It features a 4-digit back-lit display for gas concentration readings, five mode status LEDs, and a high-visibility multi-color LED light ring. Ideal for use in challenging environments, the SMART 3G-D2 offers non-intrusive calibration for precise and easy adjustments via Hall-effect switches, eliminating the need to open the instrument and declassify the area. It can reliably monitor flammable compounds (% LFL), toxic compounds (ppm), oxygen and CO2 levels (% by volume), and refrigerant gases in both ppm and % LFL. Features: • Non-intrusive one man calibration via Hall effect switches • 4-digit display and 5 mode status LEDs • Multi-color LED light ring • Standard 4-20 mA 3 wire output • Standard 3 relay output • Optional RS485 interface for Modbus communication. • Zero tracking to maintain the zero point steady from possible drifts • Ongoing system self diagnosis ATEX II 2G Ex db IIC T5 Gb (standard marking) • ATEX II 2GD Ex db IIC T5 Gb (II2 GD with optional adapter to be specified at the order) • ATEX II 2G Ex db [ib] ib IIC T6 Gb for detectors with intrinsically safe protection • SIL 2HW & SIL3 SW certified • IP rating IP65
Flammable Catalytic Sensors
| Product Code | Description | PVP | Sensor Head | Gas Cylinder
| S2157ME | Smart 3G-D2, Methane Cat 0-100% LEL | 1874.34 | Kx155ME | SIB09
| S2157GP | Smart 3G-D2, LPG Cat 0-100% LEL | 1874.34 | Kx155GP | SIB06
| S2157BU | Smart 3G-D2, Butane Cat 0-100% LEL | 1874.34 | Kx155BU | SIB06
| S2157PR | Smart 3G-D2, Propane Cat 0-100% LEL | 1874.34 | Kx155PR | SIB06
| S2157VB | Smart 3G-D2, Petrol Vapours Cat 0-100% Petrol LEL | 1874.34 | Kx155VB | SIB06
| S2157AA | Smart 3G-D2, Acetic Acid Cat 0-100% Acetic LEL | 1874.34 | Kx155AA | SIB021
| S2157AT | Smart 3G-D2, Acetone Cat 0-100% LEL | 1874.34 | Kx155AT | SIB021
| S2179AC | Smart 3G-D2, Acetylene Cat 0-100% LEL | 1874.34 | Kx151AC | SIB19
| S2157AM | Smart 3G-D2, Ammonia Cat 0-100% LEL | 1874.34 | Kx155AM | SIB06
| S2157BZ | Smart 3G-D2, Benzene Cat 0-100% LEL | 1874.34 | Kx155BZ | SIB06
| S2157CE | Smart 3G-D2, Cyclohexane Cat 0-100% LEL | 1874.34 | Kx155CE | SIB06
| S2157CP | Smart 3G-D2, Cyclopentane Cat 0-100% LEL | 1874.34 | Kx155CP | SIB06
| S2157EN | Smart 3G-D2, Ethane Cat 0-100% LEL | 1874.34 | Kx155EN | SIB06
| S2157AE | Smart 3G-D2, Ethyl Acetate Cat 0-100% Ethyl LEL | 1874.34 | Kx155AE | SIB06
| S2157ET | Smart 3G-D2, Ethyl Alcohol Cat 0-100% Ethyl LEL | 1874.34 | Kx155ET | SIB09
| S2157EE | Smart 3G-D2, Ethyl Ether Cat 0-100% Ethyl LEL | 1874.34 | Kx155EE | SIB06
| S2157EL | Smart 3G-D2, Ethylene Cat 0-100% LEL | 1874.34 | Kx155EL | SIB06
| S2157OE | Smart 3G-D2, Ethylene Oxide ETO Cat 0-100% | 1874.34 | Kx155OE | SIB06
| S2157EP | Smart 3G-D2, Heptane Cat 0-100% LEL | 1874.34 | Kx155EP | SIB06
| S2157ES | Smart 3G-D2, Hexane Cat 0-100% LEL | 1874.34 | Kx155ES | SIB06
| S2157H2 | Smart 3G-D2, Hydrogen Cat 0-100% LEL | 1874.34 | Kx155H2 | SIB19
| S2157IB | Smart 3G-D2, Iso Butane Cat 0-100% Iso LEL | 1874.34 | Kx155IB | SIB06
| S2157IT | Smart 3G-D2, Iso Butyl Alcohol Cat 0-100% Iso Butyl LEL | 1874.34 | Kx155IT | SIB021
| S2157IN | Smart 3G-D2, Iso Pentane Cat 0-100% Iso LEL | 1874.34 | Kx155IN | SIB06
| S2157IP | Smart 3G-D2, Iso Propyl Alcohol Cat 0-100% Iso Propyl LEL | 1874.34 | Kx155IP | SIB06
| S2508P8 | Smart 3G-D2, JP8 Detector, Cat 0-100% JP8 LEL | 1874.34 | Kx154P8 | SIB36
| S2157MT | Smart 3G-D2, Methyl Alcohol Cat 0-100% Methyl LEL | 1874.34 | Kx155MT | SIB09
| S2157MK | Smart 3G-D2, Methyl Ethyl Ketone MEK Cat 0-100% LEL | 1874.34 | Kx155MK | SIB06
| S2157NN | Smart 3G-D2, Nonane Cat 0-100% LEL | 1874.34 | Kx155NN | SIB021
| S2157PE | Smart 3G-D2, Pentane Cat 0-100% LEL | 1874.34 | Kx155PE | SIB06
| S2157PP | Smart 3G-D2, Propylene Cat 0-100% LEL | 1874.34 | Kx155PP | SIB09
| S2157PN | Smart 3G-D2, Propyl Alcohol Cat 0-100% LEL | 1874.34 | Kx155PN | SIB021
| S2157ST | Smart 3G-D2, Styrene Cat 0-100% LEL | 1874.34 | Kx155ST | SIB021
| S2157TM | Smart 3G-D2, Tri-Methylbenzene Cat 0-100% LEL | 1874.34 | Kx155TM | SIB06
| S2157TO | Smart 3G-D2, Toluene Cat 0-100% LEL | 1874.34 | Kx155TO | SIB06
| S2157AV | Smart 3G-D2, Vinyl Acetate Cat 0-100% Vinyl LEL | 1874.34 | Kx155AV | SIB06
| S2157XI | Smart 3G-D2, Xylene Cat 0-100% LEL | 1874.34 | Kx155XI | SIB06
Flammable Infra-Red (IR) Sensors
| Product Code | Description | PVP | Sensor Head | Gas Cylinder
| S2626ME | Smart 3G-D2, Methane IR 0-100% LEL | 2399.22 | Kx624ME | SIB09
| S2602BU | Smart 3G-D2, Butane IR 0-100% LEL | 2399.22 | Kx654BU | SIB28
| S2602PR | Smart 3G-D2, Propane IR 0-100% LEL | 2399.22 | Kx654PR | SIB28
| S2602GP | Smart 3G-D2, Lpg IR 0-100% LEL | 2399.22 | Kx654GP | SIB28
| S2602AT | Smart 3G-D2, Acetone IR 0-100% LEL | 2399.22 | Kx654AT | SIB22
| S2602BN | Smart 3G-D2, Butadiene IR 0-100% LEL | 2399.22 | Kx654BN | SIB26
| S2602CE | Smart 3G-D2, Cyclo Hexane IR 0-100% LEL | 2399.22 | Kx654CE | SIB26
| S2602DT | Smart 3G-D2, Dimethyl Ether IR 0-100% LEL | 2399.22 | Kx654DT | SIB28
| S2602EN | Smart 3G-D2, Ethane IR 0-100% LEL | 2399.22 | Kx654EN | SIB26
| S2602AE | Smart 3G-D2, Ethyl Acetate IR 0-100% Ethyl LEL | 2399.22 | Kx654AE | SIB26
| S2602ET | Smart 3G-D2, Ethyl Alcohol IR 0-100% Ethyl LEL | 2399.22 | Kx654ET | SIB28
| S2602EP | Smart 3G-D2, Heptane IR 0-100% LEL | 2399.22 | Kx654EP | SIB26
| S2602ES | Smart 3G-D2, Hexane IR 0-100% LEL | 2399.22 | Kx654ES | SIB26
| S2602IB | Smart 3G-D2, Iso Butane IR 0-100% LEL | 2399.22 | Kx654IB | SIB26
| S2602IL | Smart 3G-D2, Isobutylene IR 0-100% LEL | 2399.22 | Kx654IL | SIB26
| S2602IP | Smart 3G-D2, Iso Propyl Alcohol IR 0-100% LEL | 2399.22 | Kx654IP | SIB26
| S2602PE | Smart 3G-D2, Pentane IR 0-100% LEL | 2399.22 | Kx654PE | SIB28
| S2602PP | Smart 3G-D2, Propylene IR 0-100% LEL | 2399.22 | Kx654PP | SIB26
| S2602TO | Smart 3G-D2, Toluene IR 0-100% LEL | 2399.22 | Kx654TO | SIB22
Toxic Electrochemical Sensors
| Product Code | Description | PVP | Sensor Head | Gas Cylinder
| S2643O2 | Smart 3G-D2, Oxygen O2 EC 0-25% Vol | 2046.06 | Kx641O2 | SIB62
| S2161O2 | Smart 3G-D2, Oxygen O2 EC 0-30% Vol | 2046.06 | Kx129O2 | SIB62
| S2163CO | Smart 3G-D2, Carbon Monoxide CO EC 0-500 ppm | 2046.06 | Kx131CO | SIB56
| S2992CO | Smart 3G-D2, Carbon Monoxide CO EC 0-1000 ppm | 2046.06 | Kx586CO | SIB58
| S3227AM | Smart 3G-D2, Ammonia NH3 EC 0-100 ppm | 2656.8 | Kx225AM | RIB73-34
| S2165AM | Smart 3G-D2, Ammonia NH3 EC 0-200 ppm | 2656.8 | Kx133AM | RIB74-34
| S2167AM | Smart 3G-D2, Ammonia NH3 EC 0-1000 ppm | 2656.8 | Kx135AM | RIB76-34
| S3231AM | Smart 3G-D2, Ammonia NH3 EC 0-5000 ppm | 2656.8 | K3229AM | RIB70-34
| S2169HS | Smart 3G-D2, Hydrogen Sulphide H2S EC 0-50 ppm | 2046.06 | Kx137HS | RIB81-34
| S2873HS | Smart 3G-D2, Hydrogen Sulphide H2S EC 0-100 ppm | 2046.06 | Kx708HS | RIB81-34
| S2177NO | Smart 3G-D2, Nitric Oxide NO EC 0-100 ppm | 2656.8 | Kx145NO | RIB117-34
| S3313H2 | Smart 3G-D2, Hydrogen H2 EC 0-2000 ppm | 2656.8 | K3593H2 | RIB92-34
| S3395H2 | Smart 3G-D2, Hydrogen H2 EC 0-3000 ppm | 2656.8 | K4547H2 | RIB92-34
Intrinsically Safe version with ATEX II 2G Ex db [ib] IIC T6 Gb marking
| Product Code | Description | PVP | Sensor Head | Gas Cylinder
| S3005SD | Smart 3G-D2 IS version, Sulphur Dioxide SO2 EC 0-20 ppm | 3212.46 | K3004SD | RIB91-34
| S3009ND | Smart 3G-D2 IS version, Nitrogen Dioxide NO2 EC 0-20 ppm | 3212.46 | K3008ND | RIB116-34
| S3015CL | Smart 3G-D2 IS version, Chlorine Cl2 EC 0-20 ppm | 3212.46 | K3014CL | RIB116-34
CO2 Infra-Red (IR) Sensors
| Product Code | Description | PVP | Sensor Head | Gas Cylinder
| S3763CO2 | Smart 3G-D2, Carbon Dioxide CO2 IR 0-5000 ppm | 2410.56 | Kx444CO2 | RIB110-34
| S2634CO2 | Smart 3G-D2, Carbon Dioxide CO2 IR 0-2% Vol | 2410.56 | Kx651CO2 | SIB42
| S2635CO2 | Smart 3G-D2, Carbon Dioxide CO2 IR 0-5% Vol | 2410.56 | Kx652CO2 | SIB44
| S2449CO2 | Smart 3G-D2, Carbon Dioxide CO2 IR 0-30% Vol | 2410.56 | Kx445CO2 | SIB48
| S2636CO2 | Smart 3G-D2, Carbon Dioxide CO2 IR 0-100% Vol | 2410.56 | Kx630CO2 | SIB48
SF6 & Refrigeratent Infra-Red (IR) Sensors
| Product Code | Description | PVP | Sensor Head | Gas Cylinder
| S2602DE | Smart 3G-D2, R152a Difluoroethane IR 0-100% LEL | 2399.22 | Kx654DE | SIB26
| S4435R454B | Smart 3G-D2, R454b IR 0-100%LEL | 3398.76 | K4433R454B | SIB507
| S4829R454C | Smart 3G-D2, R454c IR 0-100%LEL | 3398.76 | K4827R454C | On Request
| S3427R32 | Smart 3G-D2, R32 IR 0-100%LEL | 3398.76 | K3425R32 | SIB512
| S3431R1234YF | Smart 3G-D2, R152a Difluoroethane IR 0-100% LEL | 3398.76 | K3429R1234YF | SIB501
| S3436R1234ZE | Smart 3G-D2, R152a Difluoroethane IR 0-100% LEL | 3398.76 | K3433R1234ZE | SIB502
| S3945R32 | Smart 3G-D2, R32 IR 0-2000 ppm | 3398.76 | K3944R32 | RIB148-34
| S3163R134A | Smart 3G-D2, R134a IR 0-2000 ppm | 3398.76 | K3596R134A | RIB144-34
| S3601SF6 | Smart 3G-D2, SF6 IR 0-2000 ppm | 3398.76 | K3599SF6 | RIB149-34
| S3661R125 | Smart 3G-D2, R125 IR 0-2000 ppm | 3398.76 | K3656R125 | RIB130-34
| S3662R1234YF | Smart 3G-D2, R1234yf IR 0-2000 ppm | 3398.76 | K3657R1234YF | RIB145-34
| S3949R1234ZE | Smart 3G-D2, R1234ze IR 0-2000 ppm | 3398.76 | K3948R1234ZE | RIB143-34
| S3663R404A | Smart 3G-D2, R404A IR 0-2000 ppm | 3398.76 | K3658R404A | RIB131-34
| S3917R407A | Smart 3G-D2, R407A IR 0-2000 ppm | 3398.76 | K3659R407A | RIB132-34
| S3664R407F | Smart 3G-D2, R407F IR 0-2000 ppm | 3398.76 | K3920R407F | RIB133-34
| S3924R449A | Smart 3G-D2, R449A IR 0-2000 ppm | 3398.76 | K3923R449A | RIB134-34
| S3665R507A | Smart 3G-D2, R507A IR 0-2000 ppm | 3398.76 | K3660R507A | RIB135-34
| S3673R410A | Smart 3G-D2, R410A IR 0-2000 ppm | 3398.76 | K3017R410A | RIB147-34
| S3953R417A | Smart 3G-D2, R417A IR 0-2000 ppm | 3398.76 | K3952R417A | RIB136-34
| S3957R448 | Smart 3G-D2, R448 IR 0-2000 ppm | 3398.76 | K3956R448 | RIB137-34
VOC Photoionisation (PID) Sensors
| Product Code | Description | PVP | Sensor Head | Gas Cylinder
| S3512IL | Smart 3G-D2, Isobutylene VOC PID 0-20 ppm | 8999.1 | On request | RIB140-34
| S3547IL | Smart 3G-D2, Isobutylene VOC PID 0-200 ppm | 8999.1 | On request | RIB141-34
| S3513IL | Smart 3G-D2, Isobutylene VOC PID 0-2000 ppm | 8999.1 | On request | RIB142-34
| S5023BZ | Smart 3G-D2, Benzene VOC PID 0-10 ppm | 8999.1 | On request | RIB140-34
| S5027BZ | Smart 3G-D2, Benzene VOC PID 0-100 ppm | 8999.1 | On request | RIB141-34
| S5031BZ | Smart 3G-D2, Benzene VOC PID 0-1000 ppm | 8999.1 | On request | RIB142-34
| S5023TO | Smart 3G-D2, Toluene VOC PID 0-10 ppm | 8999.1 | On request | RIB140-34
| S5027TO | Smart 3G-D2, Toluene VOC PID 0-100 ppm | 8999.1 | On request | RIB141-34
| S5031TO | Smart 3G-D2, Toluene VOC PID 0-1000 ppm | 8999.1 | On request | RIB142-34
SMART3G-D2 can be supplied in a stainless steel enclosure or with remote sensor head Add the desired extension to the P/N to get the product with stainless steel housing or with remote sensor head at 5 or 25m. Example: S2157ME-SW S2157ME-R25M
Additional cost to be added to the list price for Stainless Steel enclosure or remote sensor heads:
| Product Code | Description | PVP
| -SW | Exd Stainless Steel Enclosure | 868.32
| -R25M | 25m Remote Sensor Kit | 1195.56
| -R5M | 5m Remote Sensor Kit | 1138.86
| -IEC | IECEX Version | n/a
| -GD | II 2 GD Marking | 103.68
Accessories
| Product Code | Description | PVP
| STS/CSW+ | Calibration software for SMART3G, SMART3NC & SMART3R detectors consisting of a USB key, a USB/RS485 converter and the STS/IDI card. The software runs on Windows XP, Windows 7 e Windows 10. The calibration software allows to modify the alarm thresholds and the measuring range, and allows adjusting the Zero, Span and 4-20 mA values via a user-friendly menu, either in situ or in laboratory. | 626.94
| STS/IDI | Plug-in interface for RS485 serial communication suitable for SMART3G | 84.24
| STGD/AD3 | Adapter for all sensor heads (exception made for S2096__ & S2097__) to upgrade the protection from G to GD | 103.68
| RE21-PNA1KONK | ATEX cable gland for non armoured cables | 74.52

## SMART 3G-C2-LD
Auto+Hide+Values
SMART 3G-C2-LD With LED for Zone 1 category 2
The SMART 3G-C2-LD is designed to meet the demands of industrial applications, providing reliable monitoring of toxic, flammable, and refrigerant gases in classified areas. Its multi-color LED light ring offers immediate visual indication of the detector's status. Ideal for industrial environments, the SMART 3G-C2-LD features one-man calibration using the STS/CKD+ calibration keypad. These detectors can accurately monitor flammable compounds (% LFL), toxic compounds and refrigerant gases in ppm, as well as oxygen and CO2 levels (% by volume). The detectors are rated IP65 for protection against dust and water. Features: • Standard 4-20 mA 3 wire output • Standard 3 relay output • Standard multi-color LED light ring • Optional RS485 interface for Modbus communication. • Zero tracking to maintain the zero point steady from possible drifts • Ongoing system self diagnosis • ATEX II 2G Ex db IIC T5 Gb (standard marking) • ATEX II 2GD Ex db IIC T5 Gb (with optional adapter to be specified at the order) • ATEX II 2G Ex db [ib] ib IIC T6 Gb for detectors with intrinsically safe protection • SIL 2HW & SIL3 SW certified
Flammable Catalytic Sensors
| Product Code | Description | PVP | Sensor Head | Gas Cylinder
| S2155ME-LD | Smart 3G-C2-LD, Methane Cat 0-100% LEL | 1666.98 | Kx097ME | SIB09
| S2155GP-LD | Smart 3G-C2-LD, LPG Cat 0-100% LEL | 1666.98 | Kx097GP | SIB06
| S2155BU-LD | Smart 3G-C2-LD, Butane Cat 0-100% LEL | 1666.98 | Kx097BU | SIB06
| S2155PR-LD | Smart 3G-C2-LD, Propane Cat 0-100% LEL | 1666.98 | Kx097PR | SIB06
| S2155VB-LD | Smart 3G-C2-LD, Petrol Vapours Cat 0-100% LEL | 1666.98 | Kx097VB | SIB06
| S2155AA-LD | Smart 3G-C2-LD, Acetic Acid Cat 0-100% LEL | 1666.98 | Kx097AA | SIB021
| S2155AT-LD | Smart 3G-C2-LD, Acetone Cat 0-100% LEL | 1666.98 | Kx097AT | SIB021
| S2151AC-LD | Smart 3G-C2-LD, Butyl Acetate Cat 0-100% LEL | 1666.98 | Kx097AB | SIB19
| S2155AM-LD | Smart 3G-C2-LD, Butyl Alcohol Cat 0-100% LEL | 1666.98 | Kx097BT | SIB06
| S2155BZ-LD | Smart 3G-C2-LD, Cyclohexane Cat 0-100% LEL | 1666.98 | Kx097CE | SIB06
| S2155AB-LD | Smart 3G-C2-LD, Cyclopentane Cat 0-100% LEL | 1666.98 | Kx097CP | SIB021
| S2155BT-LD | Smart 3G-C2-LD, Ethane Cat 0-100% LEL | 1666.98 | Kx097EN | SIB021
| S2155CE-LD | Smart 3G-C2-LD, Ethyl Acetate Cat 0-100% LEL | 1666.98 | Kx097AE | SIB06
| S2155CP-LD | Smart 3G-C2-LD, Ethyl Alcohol Cat 0-100% LEL | 1666.98 | Kx097ET | SIB28
| S2155EN-LD | Smart 3G-C2-LD, Ethyl Ether Cat 0-100% LEL | 1666.98 | Kx097EE | SIB06
| S2155AE-LD | Smart 3G-C2-LD, Ethylene Cat 0-100% LEL | 1666.98 | Kx097EL | SIB06
| S2155ET-LD | Smart 3G-C2-LD, Heptane Cat 0-100% LEL | 1666.98 | Kx097EP | SIB09
| S2155EE-LD | Smart 3G-C2-LD, Hexane Cat 0-100% LEL | 1666.98 | Kx097ES | SIB06
| S2155EL-LD | Smart 3G-C2-LD, Hydrogen Cat 0-100% LEL | 1666.98 | Kx097H2 | SIB06
| S2155OE-LD | Smart 3G-C2-LD, Iso Butane Cat 0-100% LEL | 1666.98 | Kx097IB | SIB06
| S2155EP-LD | Smart 3G-C2-LD, Iso Butyl Alcohol Cat 0-100% LEL | 1666.98 | Kx097IT | SIB06
| S2155ES-LD | Smart 3G-C2-LD, Iso Pentane Cat 0-100% LEL | 1666.98 | Kx097IN | SIB06
| S2155H2-LD | Smart 3G-C2-LD, Iso Propyl Alcohol Cat 0-100% LEL | 1666.98 | Kx097IP | SIB19
| S2155IB-LD | Smart 3G-C2-LD, Methyl Alcohol Cat 0-100% LEL | 1666.98 | Kx097MT | SIB06
| S2155IT-LD | Smart 3G-C2-LD, Methyl Ethyl Ketone Cat 0-100% LEL | 1666.98 | Kx097MK | SIB021
| S2155IN-LD | Smart 3G-C2-LD, Nonane Cat 0-100% LEL | 1666.98 | Kx097NN | SIB06
| S2155IP-LD | Smart 3G-C2-LD, Pentane Cat 0-100% LEL | 1666.98 | Kx097PE | SIB06
| S2154P8-LD | Smart 3G-C2-LD, Propylene Cat 0-100% LEL | 1666.98 | Kx097PP | SIB36
| S2155MT-LD | Smart 3G-C2-LD, Propyl Alcohol Cat 0-100% LEL | 1666.98 | Kx097PN | SIB09
| S2155MK-LD | Smart 3G-C2-LD, Tri-Methylbenzene Cat 0-100% LEL | 1666.98 | Kx097TM | SIB06
| S2155NN-LD | Smart 3G-C2-LD, Toluene Cat 0-100% LEL | 1666.98 | Kx097TO | SIB021
| S2155PE-LD | Smart 3G-C2-LD, Xylene Cat 0-100% LEL | 1666.98 | Kx097XI | SIB06
| S2155PP-LD | Smart 3G-C2-LD, Acetylene Cat 0-100% LEL | 1666.98 | Kx151AC | SIB09
| S2155PN-LD | Smart 3G-C2-LD, Ammonia Cat 0-100% LEL | 1666.98 | Kx155AM | SIB021
| S2155ST-LD | Smart 3G-C2-LD, Benzene Cat 0-100% LEL | 1666.98 | Kx155BZ | SIB021
| S2155TM-LD | Smart 3G-C2-LD, Ethylene Oxide ETO Cat 0-100% LEL | 1666.98 | Kx155OE | SIB06
| S2155TO-LD | Smart 3G-C2-LD, Jp8 Jet Fuel Cat 0-100% LEL | 1666.98 | Kx154P8 | SIB06
| S2155XI-LD | Smart 3G-C2-LD, Styrene Cat 0-100% LEL | 1666.98 | Kx155ST | SIB06
| S2155AV-LD | Smart 3G-C2-LD, Vinyl Acetate Cat 0-100% LEL | 1666.98 | Kx155AV | SIB06
Flammable Infra-Red (IR) Sensors
| Product Code | Description | PVP | Sensor Head | Gas Cylinder
| S2624ME-LD | Smart 3G-C2-LD, Methane IR 0-100% LEL | 2182.14 | Kx624ME | SIB09
| S2654BU-LD | Smart 3G-C2-LD, Butane IR 0-100% LEL | 2182.14 | Kx654BU | SIB28
| S2654PR-LD | Smart 3G-C2-LD, Propane IR 0-100% LEL | 2182.14 | Kx654PR | SIB28
| S2654GP-LD | Smart 3G-C2-LD, LPG IR 0-100% LEL | 2182.14 | Kx654GP | SIB28
| S2654AT-LD | Smart 3G-C2-LD, Acetone IR 0-100% LEL | 2182.14 | Kx654AT | SIB22
| S2654BN-LD | Smart 3G-C2-LD, Butadiene IR 0-100% LEL | 2182.14 | Kx654BN | SIB26
| S2654CE-LD | Smart 3G-C2-LD, Cyclo Hexane IR 0-100% LEL | 2182.14 | Kx654CE | SIB26
| S2654DT-LD | Smart 3G-C2-LD, Dimethyl Ether IR 0-100% LEL | 2182.14 | Kx654DT | SIB28
| S2654EN-LD | Smart 3G-C2-LD, Ethane IR 0-100% LEL | 2182.14 | Kx654EN | SIB28
| S2654AE-LD | Smart 3G-C2-LD, Ethyl Acetate IR 0-100% LEL | 2182.14 | Kx654AE | SIB26
| S2654ET-LD | Smart 3G-C2-LD, Ethyl Alcohol IR 0-100% LEL | 2182.14 | Kx654ET | SIB28
| S2654EP-LD | Smart 3G-C2-LD, Heptane IR 0-100% LEL | 2182.14 | Kx654EP | SIB26
| S2654ES-LD | Smart 3G-C2-LD, Hexane IR 0-100% LEL | 2182.14 | Kx654ES | SIB26
| S2654IB-LD | Smart 3G-C2-LD, Iso Butane IR 0-100% LEL | 2182.14 | Kx654IB | SIB26
| S2654IL-LD | Smart 3G-C2-LD, Isobutylene IR 0-100% LEL | 2182.14 | Kx654IL | SIB26
| S2654IP-LD | Smart 3G-C2-LD, Iso Propyl Alcohol IR 0-100% LEL | 2182.14 | Kx654IP | SIB26
| S2654PE-LD | Smart 3G-C2-LD, Pentane IR 0-100% LEL | 2182.14 | Kx654PE | SIB28
| S2654PP-LD | Smart 3G-C2-LD, Propylene IR 0-100% LEL | 2182.14 | Kx654PP | SIB26
| S2654TO-LD | Smart 3G-C2-LD, Toluene IR 0-100% LEL | 2182.14 | Kx654TO | SIB22
Toxic Electrochemical Sensors and Oxygen
| Product Code | Description | PVP | Sensor Head | Gas Cylinder
| S2641O2-LD | Smart 3G-C2 LD, Oxygen O2 EC 0-25% Vol | 1731.78 | Kx641O2 | SIB62
| S2131CO-LD | Smart 3G-C2 LD, Carbon Monoxide CO EC 0-500 ppm | 1731.78 | Kx131CO | SIB56
| S2586CO-LD | Smart 3G-C2 LD, Carbon Monoxide CO EC 0-1000 ppm | 1731.78 | Kx586CO | SIB58
| S3225AM-LD | Smart 3G-C2 LD, Ammonia NH3 EC 0-100 ppm | 2357.1 | Kx225AM | RIB73-34
| S2133AM-LD | Smart 3G-C2 LD, Ammonia NH3 EC 0-200 ppm | 2357.1 | Kx133AM | RIB74-34
| S2135AM-LD | Smart 3G-C2 LD, Ammonia NH3 EC 0-1000 ppm | 2357.1 | Kx135AM | RIB76-34
| S3229AM-LD | Smart 3G-C2 LD, Ammonia NH3 EC 0-5000 ppm | 2357.1 | K3229AM | RIB70-34
| S2137HS-LD | Smart 3G-C2 LD, Hydrogen Sulphide H2S EC 0-50 ppm | 1731.78 | Kx137HS | RIB81-34
| S3708HS-LD | Smart 3G-C2 LD, Hydrogen Sulphide H2S EC 0-100 ppm | 1731.78 | Kx708HS | RIB81-34
| S2145NO-LD | Smart 3G-C2 LD, Nitric Oxide NO EC 0-100 ppm | 1731.78 | Kx145NO | RIB117-34
| S3593H2-LD | Smart 3G-C2 LD, Hydrogen H2 EC 0-2000 ppm | 2357.1 | K3593H2 | RIB92-34
| S4547H2-LD | Smart 3G-C2 LD, Hydrogen H2 EC 0-3000 ppm | 2357.1 | K4547H2 | RIB92-34
Intrinsically Safe Sensor Head. ATEX II2G Ex db [ib] ib IIC T6 Gb
| Product Code | Description | PVP | Sensor Head | Gas Cylinder
| S3004SD-LD | Smart 3G-C2 LD, Sulphur Dioxide SO2 EC 0-20 ppm | 3000.24 | K3004SD | RIB91-34
| S3008ND-LD | Smart 3G-C2 LD, Nitrogen Dioxide NO2 EC 0-20 ppm | 3000.24 | K3008ND | RIB116-34
| S3014CL-LD | Smart 3G-C2 LD, Chlorine CL2 EC 0-20 ppm | 3000.24 | K3014CL | RIB116-34
CO2 Infra-Red (IR) Sensors
| Product Code | Description | PVP | Sensor Head | Gas Cylinder
| S3893CO2-LD | Smart 3G-C2 LD, Carbon Dioxide CO2 IR 0-5000 ppm | 2191.86 | Kx444CO2 | RIB110-34
| S2651CO2-LD | Smart 3G-C2 LD, Carbon Dioxide CO2 IR 0-2% Vol | 2191.86 | Kx651CO2 | SIB42
| S2652CO2-LD | Smart 3G-C2 LD, Carbon Dioxide CO2 IR 0-5% Vol | 2191.86 | Kx652CO2 | SIB44
| S2445CO2-LD | Smart 3G-C2 LD, Carbon Dioxide CO2 IR 0-30% Vol | 2191.86 | Kx445CO2 | SIB48
| S2630CO2-LD | Smart 3G-C2 LD, Carbon Dioxide CO2 IR 0-100% Vol | 2191.86 | Kx630CO2 | SIB48
Infra-Red (IR) Sensors - Refrigerant Gases A1, A2L and SF6
| Product Code | Description | PVP | Sensor Head | Gas Cylinder
| S2654DE-LD | Smart 3G-C2 LD, R152a IR 0-100% LEL | 2182.14 | Kx654DE | SIB26
| S4433R454B-LD | Smart 3G-C2 LD, R454b IR 0-100% LEL | 3278.88 | K4433R454B | SIB507
| S4827R454C-LD | Smart 3G-C2 LD, R454c IR 0-100% LEL | 3278.88 | K4827R454C | On Request
| S3425R32-LD | Smart 3G-C2 LD, R32 IR 0-100% LEL | 3278.88 | K3425R32 | SIB512
| S3429R1234YF-LD | Smart 3G-C2 LD, R1234yf IR 0-100% LEL | 3278.88 | K3429R1234YF | SIB501
| S3433R1234ZE-LD | Smart 3G-C2 LD, R1234ze IR 0-100% LEL | 3278.88 | K3433R1234ZE | SIB502
| S3944R32-LD | Smart 3G-C2 LD, R32 IR 0-2000 ppm | 3278.88 | K3944R32 | RIB148-34
| S3596R134A-LD | Smart 3G-C2 LD, R134a IR 0-2000 ppm | 3278.88 | K3596R134A | RIB144-34
| S3599SF6-LD | Smart 3G-C2 LD, SF6 IR 0-2000 ppm | 3278.88 | K3599SF6 | RIB149-34
| S3656R125-LD | Smart 3G-C2 LD, R125 IR 0-2000 ppm | 3278.88 | K3656R125 | RIB130-34
| S3657R1234YF-LD | Smart 3G-C2 LD, R1234yf IR 0-2000 ppm | 3278.88 | K3657R1234YF | RIB145-34
| S3948R1234ZE-LD | Smart 3G-C2 LD, R1234ze IR 0-2000 ppm | 3278.88 | K3948R1234ZE | RIB143-34
| S3658R404A-LD | Smart 3G-C2 LD, R404A IR 0-2000 ppm | 3278.88 | K3658R404A | RIB131-34
| S3659R407A-LD | Smart 3G-C2 LD, R407A IR 0-2000 ppm | 3278.88 | K3659R407A | RIB132-34
| S3920R407F-LD | Smart 3G-C2 LD, R407F IR 0-2000 ppm | 3278.88 | K3920R407F | RIB133-34
| S3923R449A-LD | Smart 3G-C2 LD, R449 IR 0-2000 ppm | 3278.88 | K3923R449A | RIB134-34
| S3660R507A-LD | Smart 3G-C2 LD, R507A IR 0-2000 ppm | 3278.88 | K3660R507A | RIB135-34
| S3017R410A-LD | Smart 3G-C2 LD, R410A IR 0-2000 ppm | 3278.88 | K3017R410A | RIB147-34
| S3952R417A-LD | Smart 3G-C2 LD, R417A IR 0-2000 ppm | 3278.88 | K3952R417A | RIB136-34
| S3956R448-LD | Smart 3G-C2 LD, R448 IR 0-2000 ppm | 3278.88 | K3956R448 | RIB137-34
VOC Photoionisation (PID) Sensors
| Product Code | Description | PVP | Sensor Head | Gas Cylinder
| S3541IL-LD | Smart 3G-C2 LD, Isobutylene VOC PID 0-20 ppm | 8822.52 | On request | RIB140-34
| S3545IL-LD | Smart 3G-C2 LD, Isobutylene VOC PID 0-200 ppm | 8822.52 | On request | RIB141-34
| S5122IL-LD | Smart 3G-C2 LD, Isobutylene VOC PID 0-2000 ppm | 8822.52 | On request | RIB142-34
| S5021BZ-LD | Smart 3G-C2 LD, Benzene VOC PID 0-10 ppm | 8822.52 | On request | RIB140-34
| S5025BZ-LD | Smart 3G-C2 LD, Benzene VOC PID 0-100 ppm | 8822.52 | On request | RIB141-34
| S5029BZ-LD | Smart 3G-C2 LD, Benzene VOC PID 0-1000 ppm | 8822.52 | On request | RIB142-34
| S5021TO-LD | Smart 3G-C2 LD, Toluene VOC PID 0-10 ppm | 8822.52 | On request | RIB140-34
| S5025TO-LD | Smart 3G-C2 LD, Toluene VOC PID 0-100 ppm | 8822.52 | On request | RIB141-34
| S5029TO-LD | Smart 3G-C2 LD, Toluene VOC PID 0-1000 ppm | 8822.52 | On request | RIB142-34
Additional cost to be added to the list price for different versions
| Product Code | Description | PVP
| -SD | Exd Stainless Steel Enclosure | 868.32
| -R25M | 25m Remote Sensor Kit (Cable NOT Included) | 1195.56
| -R5M | 5m Remote Sensor Kit (Cable Included) | 1138.86
| -IEC | IECEX Version | 0
| -GD | II 2 GD Marking | 103.68
Accessories
| Product Code | Description | PVP
| STS/CSW+ | Calibration software for SMART3G, SMART3NC & SMART3R detectors consisting of a USB key, a USB/RS485 converter and the STS/IDI card. The software runs on Windows XP, Windows 7 e Windows 10. The calibration software allows to modify the alarm thresholds and the measuring range, and allows adjusting the Zero, Span and 4-20 mA values via a user-friendly menu, either in situ or in laboratory. | 626.94
| STS/IDI | Plug-in interface for RS485 serial communication suitable for SMART3G | 84.24
| STS/CKD+ | Enhanced calibration keypad with 7-digit display. It can be connected to SENSITRON’S SMART detectors to adjust the Zero, Span and 4-20 mA values. | 515.16
| STS/CKD-OLED | Enhanced calibration keypad with OLED display. It can be connected to all SENSITRON’S SMART detectors to adjust the Zero, Span and 4-20 mA values. This version offers the visualization of the instantaneous concentration in graphic format. | 780.84
| STGD/AD3 | Adapter for all sensor heads (exception made for S2096__ & S2097__) to upgrade the protection from G to GD | 103.68
| RE21-PNA1KONK | ATEX cable gland for non armoured cables | 74.52

## SMART 3G-C2
Auto+Hide+Values
SMART 3G-C2 For Zone 1 Category 2
The SMART 3G-C2 is designed to meet industrial requirements, enabling the monitoring of toxic, flammable, and refrigerant gases in industrial environments and classified areas. Ideal for industrial settings, the SMART 3G-C2 features one-man calibration using the STS/CKD+ calibration keypad. These detectors can reliably monitor flammable compounds (% LFL), toxic compounds and refrigerant gases in ppm, as well as oxygen and CO2 levels (% by volume). The detectors are rated IP65 for protection against dust and water. Features: • Standard 4-20 mA 3 wire output • Optional 1 or 3 relay output • Optional RS485 interface for Modbus communication. • Zero tracking to maintain the zero point steady from possible drifts • Ongoing system self diagnosis • ATEX II 2G Ex db IIC T5 Gb (standard marking) (II 2GD with optional adapter) • ATEX II 2G Ex db [ib] ib IIC T6 Gb for detectors with intrinsically safe protection • SIL 2 HW & SIL3 SW certified
Flammable Catalytic Sensors
| Product Code | Description | PVP | Sensor Head |  | Gas Cylinder
| S2097ME | Smart 3G-C2, Methane Cat 0-100% LEL | 1023.84 | Kx097ME |  | SIB09
| S2097GP | Smart 3G-C2, LPG Cat 0-100% LEL | 1023.84 | Kx097GP | #REF! | SIB06
| S2097BU | Smart 3G-C2, Butane Cat 0-100% LEL | 1023.84 | Kx097BU | #REF! | SIB06
| S2097PR | Smart 3G-C2, Propane Cat 0-100% LEL | 1023.84 | Kx097PR | #REF! | SIB09
| S2097VB | Smart 3G-C2, Petrol Vapours Cat 0-100% LEL | 1023.84 | Kx097VB | #REF! | SIB06
| S2097AA | Smart 3G-C2, Acetic Acid Cat 0-100% LEL | 1023.84 | Kx097AA | #REF! | SIB021
| S2097AT | Smart 3G-C2, Acetone Cat 0-100% LEL | 1023.84 | Kx097AT | #REF! | SIB06
| S2097AB | Smart 3G-C2, Butyl Acetate Cat 0-100% LEL | 1023.84 | Kx097AB | #REF! | SIB06
| S2097BT | Smart 3G-C2, Butyl Alcohol Cat 0-100% LEL | 1023.84 | Kx097BT | #REF! | SIB06
| S2097CE | Smart 3G-C2, Cyclohexane Cat 0-100% LEL | 1023.84 | Kx097CE | #REF! | SIB06
| S2097CP | Smart 3G-C2, Cyclopentane Cat 0-100% LEL | 1023.84 | Kx097CP | #REF! | SIB06
| S2097EN | Smart 3G-C2, Ethane Cat 0-100% LEL | 1023.84 | Kx097EN | #REF! | SIB09
| S2097AE | Smart 3G-C2, Ethyl Acetate Cat 0-100% LEL | 1023.84 | Kx097AE | #REF! | SIB06
| S2097ET | Smart 3G-C2, Ethyl Alcohol Cat 0-100% LEL | 1023.84 | Kx097ET | #REF! | SIB06
| S2097EE | Smart 3G-C2, Ethyl Ether Cat 0-100% LEL | 1023.84 | Kx097EE | #REF! | SIB06
| S2097EL | Smart 3G-C2, Ethylene Cat 0-100% LEL | 1023.84 | Kx097EL | #REF! | SIB06
| S2097EP | Smart 3G-C2, Heptane Cat 0-100% LEL | 1023.84 | Kx097EP | #REF! | SIB06
| S2097ES | Smart 3G-C2, Hexane Cat 0-100% LEL | 1023.84 | Kx097ES | #REF! | SIB06
| S2097H2 | Smart 3G-C2, Hydrogen Cat 0-100% LEL | 1023.84 | Kx097H2 | #REF! | SIB19
| S2097IB | Smart 3G-C2, Iso Butane Cat 0-100% LEL | 1023.84 | Kx097IB | #REF! | SIB09
| S2097IT | Smart 3G-C2, Iso Butyl Alcohol Cat 0-100% LEL | 1023.84 | Kx097IT | #REF! | SIB021
| S2097IN | Smart 3G-C2, Iso Pentane Cat 0-100% LEL | 1023.84 | Kx097IN | #REF! | SIB06
| S2097IP | Smart 3G-C2, Iso Propyl Alcohol Cat 0-100% LEL | 1023.84 | Kx097IP | #REF! | SIB06
| S2097MT | Smart 3G-C2, Methyl Alcohol Cat 0-100% LEL | 1023.84 | Kx097MT | #REF! | SIB09
| S2097MK | Smart 3G-C2, Methyl Ethyl Ketone Cat 0-100% LEL | 1023.84 | Kx097MK | #REF! | SIB06
| S2097NN | Smart 3G-C2, Nonane Cat 0-100% LEL | 1023.84 | Kx097NN | #REF! | SIB021
| S2097PE | Smart 3G-C2, Pentane Cat 0-100% LEL | 1023.84 | Kx097PE | #REF! | SIB06
| S2097PP | Smart 3G-C2, Propylene Cat 0-100% LEL | 1023.84 | Kx097PP | #REF! | SIB09
| S2097PN | Smart 3G-C2, Propyl Alcohol Cat 0-100% LEL | 1023.84 | Kx097PN | #REF! | SIB06
| S2097TM | Smart 3G-C2, Tri-Methylbenzene Cat 0-100% LEL | 1023.84 | Kx097TM | #REF! | SIB09
| S2097TO | Smart 3G-C2, Toluene Cat 0-100% LEL | 1023.84 | Kx097TO | #REF! | SIB06
| S2097XI | Smart 3G-C2, Xylene Cat 0-100% LEL | 1023.84 | Kx097XI | #REF! | SIB06
| S2151AC | Smart 3G-C2, Acetylene Cat 0-100% LEL | 1155.06 | Kx151AC | #REF! | SIB19
| S2155AM | Smart 3G-C2, Ammonia Cat 0-100% LEL | 1155.06 | Kx155AM | #REF! | SIB06
| S2155BZ | Smart 3G-C2, Benzene Cat 0-100% LEL | 1155.06 | Kx155BZ | #REF! | SIB06
| S2155OE | Smart 3G-C2, Ethylene Oxide ETO Cat 0-100% LEL | 1155.06 | Kx155OE | #REF! | SIB06
| S2154P8 | Smart 3G-C2, Jp8 Jet Fuel Cat 0-100% LEL | 1221.48 | Kx154P8 | #REF! | SIB36
| S2155ST | Smart 3G-C2, Styrene Cat 0-100% LEL | 1155.06 | Kx155ST | #REF! | SIB021
| S2155AV | Smart 3G-C2, Vinyl Acetate Cat 0-100% LEL | 1155.06 | Kx155AV | #REF! | SIB06
Flammable Infra-Red (IR) Sensors
| Product Code | Description | PVP | Sensor Head |  | Gas Cylinder
| S2624ME | Smart 3G-C2, Methane IR 0-100% LEL | 1874.34 | Kx624ME |  | SIB09
| S2654BU | Smart 3G-C2, Butane IR 0-100% LEL | 1874.34 | Kx654BU | #REF! | SIB28
| S2654PR | Smart 3G-C2, Propane IR 0-100% LEL | 1874.34 | Kx654PR | #REF! | SIB28
| S2654GP | Smart 3G-C2, LPG IR 0-100% LEL | 1874.34 | Kx654GP | #REF! | SIB28
| S2654AT | Smart 3G-C2, Acetone IR 0-100% LEL | 1874.34 | Kx654AT | #REF! | SIB22
| S2654BN | Smart 3G-C2, Butadiene IR 0-100% LEL | 1874.34 | Kx654BN | #REF! | SIB26
| S2654CE | Smart 3G-C2, Cyclo Hexane IR 0-100% LEL | 1874.34 | Kx654CE | #REF! | SIB26
| S2654DT | Smart 3G-C2, Dimethyl Ether IR 0-100% LEL | 1874.34 | Kx654DT | #REF! | SIB28
| S2654EN | Smart 3G-C2, Ethane IR 0-100% LEL | 1874.34 | Kx654EN | #REF! | SIB28
| S2654AE | Smart 3G-C2, Ethyl Acetate IR 0-100% LEL | 1874.34 | Kx654AE | #REF! | SIB26
| S2654ET | Smart 3G-C2, Ethyl Alcohol IR 0-100% LEL | 1874.34 | Kx654ET | #REF! | SIB28
| S2654EP | Smart 3G-C2, Heptane IR 0-100% LEL | 1874.34 | Kx654EP | #REF! | SIB26
| S2654ES | Smart 3G-C2, Hexane IR 0-100% LEL | 1874.34 | Kx654ES | #REF! | SIB26
| S2654IB | Smart 3G-C2, Iso Butane IR 0-100% LEL | 1874.34 | Kx654IB | #REF! | SIB26
| S2654IL | Smart 3G-C2, Isobutylene IR 0-100% LEL | 1874.34 | Kx654IL | #REF! | SIB26
| S2654IP | Smart 3G-C2, Iso Propyl Alcohol IR 0-100% LEL | 1874.34 | Kx654IP | #REF! | SIB26
| S2654PE | Smart 3G-C2, Pentane IR 0-100% LEL | 1874.34 | Kx654PE | #REF! | SIB28
| S2654PP | Smart 3G-C2, Propylene IR 0-100% LEL | 1874.34 | Kx654PP | #REF! | SIB26
| S2654TO | Smart 3G-C2,Toluene IR 0-100% LEL | 1874.34 | Kx654TO | #REF! | SIB22
Toxic Electrochemical Sensors and Oxygen
| Product Code | Description | PVP | Sensor Head |  | Gas Cylinder
| S2641O2 | Smart 3G-C2, Oxygen O2 EC 0-25% Vol | 1432.08 | Kx641O2 |  | SIB62
| S2129O2 | Smart 3G-C2, Oxygen O2 EC 0-30% Vol | 1432.08 | Kx129O2 |  | SIB62
| S2131CO | Smart 3G-C2, Carbon Monoxide CO EC 0-500 ppm | 1432.08 | Kx131CO | #REF! | SIB56
| S2586CO | Smart 3G-C2, Carbon Monoxide CO EC 0-1000 ppm | 1432.08 | Kx586CO | #REF! | SIB58
| S3225AM | Smart 3G-C2, Ammonia NH3 EC 0-100 ppm | 2046.06 | Kx225AM | #REF! | RIB73-34
| S2133AM | Smart 3G-C2, Ammonia NH3 EC 0-200 ppm | 2046.06 | Kx133AM | #REF! | RIB74-34
| S2135AM | Smart 3G-C2, Ammonia NH3 EC 0-1000 ppm | 2046.06 | Kx135AM | #REF! | RIB76-34
| S3229AM | Smart 3G-C2, Ammonia NH3 EC 0-5000 ppm | 2046.06 | K3229AM | #REF! | RIB70-34
| S2137HS | Smart 3G-C2, Hydrogen Sulphide H2S EC 0-50 ppm | 1432.08 | Kx137HS | #REF! | RIB81-34
| S3708HS | Smart 3G-C2, Hydrogen Sulphide H2S EC 0-100 ppm | 1432.08 | Kx708HS | #REF! | RIB81-34
| S2145NO | Smart 3G-C2, Nitric Oxide NO EC 0-100 ppm | 2046.06 | Kx145NO | #REF! | RIB117-34
| S3593H2 | Smart 3G-C2, Hydrogen H2 EC 0-2000 ppm | 2046.06 | K3593H2 | #REF! | RIB92-34
| S4547H2 | Smart 3G-C2, Hydrogen H2 EC 0-3000 ppm | 2046.06 | K4547H2 | #REF! | RIB92-34
Detectors with intrinsically safe head. ATEX II2G Ex db [ib] ib IIC T6 Gb
| Product Code | Description | PVP | Sensor Head |  | Gas Cylinder
| S3004SD | Smart 3G-C2, Sulphur Dioxide SO2 EC 0-20 ppm | 2784.78 | K3004SD | #REF! | RIB91-34
| S3008ND | Smart 3G-C2, Nitrogen Dioxide NO2 EC 0-20 ppm | 2784.78 | K3008ND | #REF! | RIB116-34
| S3014CL | Smart 3G-C2, Chlorine CL2 EC 0-20 ppm | 2784.78 | K3014CL | #REF! | RIB116-34
CO2 Infra-Red (IR) Sensors
| Product Code | Description | PVP | Sensor Head |  | Gas Cylinder
| S3893CO2 | Smart 3G-C2, Carbon Dioxide CO2 IR 0-5000 ppm | 1884.06 | Kx444CO2 |  | RIB110-34
| S2651CO2 | Smart 3G-C2, Carbon Dioxide CO2 IR 0-2% Vol | 1884.06 | Kx651CO2 | #REF! | SIB42
| S2652CO2 | Smart 3G-C2, Carbon Dioxide CO2 IR 0-5% Vol | 1884.06 | Kx652CO2 | #REF! | SIB44
| S2445CO2 | Smart 3G-C2, Carbon Dioxide CO2 IR 0-30% Vol | 1884.06 | Kx445CO2 | #REF! | SIB48
| S2630CO2 | Smart 3G-C2, Carbon Dioxide CO2 IR 0-100% Vol | 1884.06 | Kx630CO2 | #REF! | SIB48
Infra-Red (IR) Sensors - Refrigerant Gases A1, A2L and SF6
| Product Code | Description | PVP | Sensor Head |  | Gas Cylinder
| S2654DE | Smart 3G-C2, R152a IR 0-100% LEL | 1874.34 | Kx654DE |  | SIB26
| S4433R454B | Smart 3G-C2, R454b IR 0-100% LEL | 3095.82 | K4433R454B |  | SIB507
| S4827R454C | Smart 3G-C2, R454c IR 0-100% LEL | 3095.82 | K4827R454C | #REF! | On Request
| S3425R32 | Smart 3G-C2, R32 IR 0-100% LEL | 3095.82 | K3425R32 | #REF! | SIB512
| S3429R1234YF | Smart 3G-C2, R1234yf IR 0-100% LEL | 3095.82 | K3429R1234YF | #REF! | SIB501
| S3433R1234ZE | Smart 3G-C2, R1234ze IR 0-100% LEL | 3095.82 | K3433R1234ZE | #REF! | SIB502
| S3944R32 | Smart 3G-C2, R32 IR 0-2000 ppm | 3095.82 | K3944R32 | #REF! | RIB148-34
| S3596R134A | Smart 3G-C2, R134a IR 0-2000 ppm | 3095.82 | K3596R134A | #REF! | RIB144-34
| S3599SF6 | Smart 3G-C2, SF6 IR 0-2000 ppm | 3095.82 | K3599SF6 | #REF! | RIB149-34
| S3656R125 | Smart 3G-C2, R125 IR 0-2000 ppm | 3095.82 | K3656R125 | #REF! | RIB130-34
| S3657R1234YF | Smart 3G-C2, R1234yf IR 0-2000 ppm | 3095.82 | K3657R1234YF | #REF! | RIB145-34
| S3948R1234ZE | Smart 3G-C2, R1234ze IR 0-2000 ppm | 3095.82 | K3948R1234ZE | #REF! | RIB143-34
| S3658R404A | Smart 3G-C2, R404A IR 0-2000 ppm | 3095.82 | K3658R404A | #REF! | RIB131-34
| S3659R407A | Smart 3G-C2, R407A IR 0-2000 ppm | 3095.82 | K3659R407A | #REF! | RIB132-34
| S3920R407F | Smart 3G-C2, R407F IR 0-2000 ppm | 3095.82 | K3920R407F | #REF! | RIB133-34
| S3923R449A | Smart 3G-C2, R449 IR 0-2000 ppm | 3095.82 | K3923R449A | #REF! | RIB134-34
| S3660R507A | Smart 3G-C2, R507A IR 0-2000 ppm | 3095.82 | K3660R507A | #REF! | RIB135-34
| S3017R410A | Smart 3G-C2, R410A IR 0-2000 ppm | 3095.82 | K3017R410A | #REF! | RIB147-34
| S3952R417A | Smart 3G-C2, R417A IR 0-2000 ppm | 3095.82 | K3952R417A | #REF! | RIB136-34
| S3956R448 | Smart 3G-C2, R448 IR 0-2000 ppm | 3095.82 | K3956R448 | #REF! | RIB137-34
VOC Photoionisation (PID) Sensors
| Product Code | Description | PVP | Sensor Head |  | Gas Cylinder
| S3541IL | Smart 3G-C2, Isobutylene VOC PID 0-20 ppm | 8469.36 | On request |  | RIB140-34
| S3545IL | Smart 3G-C2, Isobutylene VOC PID 0-200 ppm | 8469.36 | On request |  | RIB141-34
| S5122IL | Smart 3G-C2, Isobutylene VOC PID 0-2000 ppm | 8469.36 | On request |  | RIB142-34
| S5021BZ | Smart 3G-C2, Benzene VOC PID 0-10 ppm | 8469.36 | On request |  | RIB140-34
| S5025BZ | Smart 3G-C2, Benzene VOC PID 0-100 ppm | 8469.36 | On request |  | RIB141-34
| S5029BZ | Smart 3G-C2, Benzene VOC PID 0-1000 ppm | 8469.36 | On request |  | RIB142-34
| S5021TO | Smart 3G-C2, Toluene VOC PID 0-10 ppm | 8469.36 | On request |  | RIB140-34
| S5025TO | Smart 3G-C2, Toluene VOC PID 0-100 ppm | 8469.36 | On request |  | RIB141-34
| S5029TO | Smart 3G-C2, Toluene VOC PID 0-1000 ppm | 8469.36 | On request |  | RIB142-34
SMART3G-C2 can be supplied in a stainless steel enclosure or with an Enclosure+Electronics Kit (remote head). Add the desired suffix to the P/N to identify the correct instrument: Example: S2097ME-SW S2097ME-R25M
Additional cost to be added to the list price for different versions
| Product Code | Description | PVP
| -SD | Stainless steel enclosure | 868.32
| -R25M | 25m remote sensor kit to separate the sensor head from the transmitter | 1195.56
| -R5M | 5m remote sensor kit to separate the sensor head from the transmitter | 1138.86
| -IEC | Detector with IECEX certification | 0
| -GD | For detectors marked as II 2 GD | 103.68
Accessories
| Product Code | Description | PVP
| STS/CSW+ | Calibration software for SMART3G, SMART3NC & SMART3R detectors consisting of a USB key, a USB/RS485 converter and the STS/IDI card. The software runs on Windows XP, Windows 7 e Windows 10. The calibration software allows to modify the alarm thresholds and the measuring range, and allows adjusting the Zero, Span and 4-20 mA values via a user-friendly menu, either in situ or in laboratory. | 626.94
| STS/IDI | Plug-in interface for RS485 serial communication suitable for SMART3G | 84.24
| STS/CKD+ | Enhanced calibration keypad with 7-digit display. It can be connected to SENSITRON’S SMART detectors to adjust the Zero, Span and 4-20 mA values. | 515.16
| STS/CKD-OLED | Enhanced calibration keypad with OLED display. It can be connected to all SENSITRON’S SMART detectors to adjust the Zero, Span and 4-20 mA values. This version offers the visualization of the instantaneous concentration in graphic format. | 780.84
| STS1REL | 1-Relay card for SMART3G. Relay at 12÷24V | 84.24
| STS3REL | 3-Relay card for SMART3G. Relay at 12÷24V | 124.74
| STGD/AD3 | Adapter for type 3 sensor heads to upgrade the protection from G to GD | 103.68
| STGD/AD2 | Adapter for type 2 sensor heads to upgrade the protection from G to GD | 103.68
| RE21-PNA1KONK | Cable gland for non-armoured cable, ATEX | 74.52

## SMART 3G-D3
Auto+Hide+Values
SMART 3G-D3 With display for zone 2 category 3 areas
The SMART 3G-D3 is designed to meet industrial requirements, enabling the monitoring of toxic and flammable gases in Zone 2 classified areas. It features a 4-digit back-lit display for gas concentration readings, five mode status LEDs, and a high-visibility multi-color LED light ring. The SMART 3G-D3 offers non-intrusive calibration for accurate and easy adjustments via Hall-effect switches, without the need to open the instrument and declassify the area. These detectors can reliably monitor flammable compounds (% LFL), toxic compounds in ppm, as well as oxygen and CO2 levels (% by volume). Features: • Non-intrusive one man calibration via Hall effect switches • 4-digit display and 5 mode status LEDs • Multi-color LED light ring • Standard 4-20 mA 3 wire output • Standard 3 relay output • Optional RS485 interface for Modbus communication. • Zero tracking to maintain the zero point steady from possible drifts • Ongoing system self diagnosis • ATEX: II 3G Ex db ec nC IIC T5 Gc • SIL 2 HW & SIL3 SW certified • IP55
Flammable Catalytic Sensors
| Product Code | Description | PVP | Sensor Head | Gas Cylinder
| S2156ME | Smart 3G-D3, Methane Cat 0-100% LEL | 1466.1 | Kx155ME | SIB09
| S2156GP | Smart 3G-D3, LPG Cat 0-100% LEL | 1466.1 | Kx155GP | SIB06
| S2156PR | Smart 3G-D3, Propane Cat 0-100% LEL | 1466.1 | Kx155BU | SIB06
| S2156BU | Smart 3G-D3, Butane Cat 0-100% LEL | 1466.1 | Kx155PR | SIB06
| S2156VB | Smart 3G-D3, Petrol Vapours Cat 0-100% LEL | 1466.1 | Kx155VB | SIB06
| S2156H2 | Smart 3G-3, Hydrogen Cat 0-100% LEL | 1466.1 | Kx155H2 | SIB19
Flammable Infra-Red (IR) Sensors
| Product Code | Description | PVP | Sensor Head | Gas Cylinder
| S2625ME | Smart 3G-D3, Methane IR 0-100% LEL | 1992.6 | Kx624ME | SIB09
| S2601PR | Smart 3G-D3, Propane IR 0-100% LEL | 1992.6 | Kx654PR | SIB28
| S2601BU | Smart 3G-D3, Butane IR 0-100% LEL | 1992.6 | Kx654BU | SIB28
| S2601GP | Smart 3G-D3, LPG IR 0-100% LEL | 1992.6 | Kx654GP | SIB28
Toxic Electrochemical Sensors
| Product Code | Description | PVP | Sensor Head | Gas Cylinder
| S2642O2 | Smart 3G-D3, Oxygen O2 EC 0-25% Vol | 1539 | Kx641O2 | SIB62
| S2159O2 | Smart 3G-D3, Oxygen O2 EC 0-30% Vol | 1539 | Kx129O2 | SIB62
| S2162CO | Smart 3G-D3, Carbon Monoxide CO EC 0-500 ppm | 1539 | Kx131CO | SIB56
| S3226AM | Smart 3G-D3, Ammonia NH3 EC 0-100 ppm | 2140.02 | Kx225AM | RIB73-34
| S2164AM | Smart 3G-D3, Ammonia NH3 EC 0-200 ppm | 2140.02 | Kx133AM | RIB74-34
| S2166AM | Smart 3G-D3, Ammonia NH3 EC 0-1000 ppm | 2140.02 | Kx135AM | RIB76-34
| S3230AM | Smart 3G-D3, Ammonia NH3 EC 0-5000 ppm | 2140.02 | K3229AM | RIB70-34
| S2168HS | Smart 3G-D3, Hydrogen Sulphide H2S EC 0-50 ppm | 1539 | Kx137HS | RIB81-34
| S2176NO | Smart 3G-D3, Nitric Oxide NO EC 0-100 ppm | 2140.02 | Kx145NO | RIB117-34
| S3594H2 | Smart 3G-D3, Hydrogen H2 EC 0-2000 ppm | 2140.02 | K3593H2 | RIB92-34
| S4548H2 | Smart 3G-D3, Hydrogen H2 EC 0-3000 ppm | 2140.02 | K4547H2 | RIB92-34
CO2 Infra-Red (IR) Sensors
| Product Code | Description | PVP | Sensor Head | Gas Cylinder
| S2631CO2 | Smart 3G-D3, Carbon Dioxide CO2 IR 0-2% Vol | 1992.6 | Kx651CO2 | SIB42
| S2632CO2 | Smart 3G-D3, Carbon Dioxide CO2 IR 0-5% Vol | 1992.6 | Kx652CO2 | SIB44
| S2447CO2 | Smart 3G-D3, Carbon Dioxide CO2 IR 0-30% Vol | 1992.6 | Kx445CO2 | SIB48
| S2448CO2 | Smart 3G-D3, Carbon Dioxide CO2 IR 0-5000 ppm | 1992.6 | Kx444CO2 | RIB110-34
| S2633CO2 | Smart 3G-D3, Carbon Dioxide CO2 IR 0-100% Vol | 1992.6 | Kx630CO2 | SIB48
SF6 & Refrigeratent Infra-Red (IR) Sensors
| Product Code | Description | PVP | Sensor Head | Gas Cylinder
| S4434R454B | Smart 3G-D3, R454b IR 0-100% LEL | 3262.68 | K4433R454B | SIB507
| S4828R454C | Smart 3G-D3, R454c IR 0-100% LEL | 3262.68 | K4827R454C | On Request
| S3426R32 | Smart 3G-D3, R32 IR 0-100% LEL | 3262.68 | K3425R32 | SIB512
| S3430R1234YF | Smart 3G-D3, R1234yf IR 0-100% LEL | 3262.68 | K3429R1234YF | SIB501
| S3435R1234ZE | Smart 3G-D3, R1234ze IR 0-100% LEL | 3262.68 | K3433R1234ZE | SIB502
| S3943R32 | Smart 3G-D3, R32 IR 0-2000 ppm | 3262.68 | Kx944R32 | RIB148-34
| S3597R134A | Smart 3G-D3, R134a IR 0-2000 ppm | 3262.68 | K3596R134A | RIB144-34
| S3600SF6 | Smart 3G-D3, SF6 IR 0-2000 ppm | 3262.68 | K3599SF6 | RIB149-34
| S3651R125 | Smart 3G-D3, R125 IR 0-2000 ppm | 3262.68 | K3656R125 | RIB130-34
| S3652R1234YF | Smart 3G-D3, R1234yf IR 0-2000 ppm | 3262.68 | K3657R1234YF | RIB145-34
| S3947R1234ZE | Smart 3G-D3, R1234ze IR 0-2000 ppm | 3262.68 | K3948R1234ZE | RIB143-34
| S3653R404A | Smart 3G-D3, R404a IR 0-2000 ppm | 3262.68 | K3658R404A | RIB131-34
| S3654R407A | Smart 3G-D3, R407a IR 0-2000 ppm | 3262.68 | K3659R407A | RIB132-34
| S3919R407F | Smart 3G-D3, R407f IR 0-2000 ppm | 3262.68 | K3920R407F | RIB133-34
| S3922R449A | Smart 3G-D3, R449 IR 0-2000 ppm | 3262.68 | K3923R449A | RIB134-34
| S3655R507A | Smart 3G-D3, R507a IR 0-2000 ppm | 3262.68 | K3660R507A | RIB135-34
| S3672R410A | Smart 3G-D3, R410a IR 0-2000 ppm | 3262.68 | K3017R410A | RIB147-34
| S3951R417A | Smart 3G-D3, R417a IR 0-2000 ppm | 3262.68 | K3952R417A | RIB136-34
| S3955R448 | Smart 3G-D3, R448 IR 0-2000 ppm | 3262.68 | K3956R448 | RIB137-34
SMART3G-D3 - VOC detectors (PID sensors)
| Product Code | Description | PVP | Sensor Head | Gas Cylinder
| S3542IL | Smart 3G-D3, Isobutylene VOC PID 0-20 ppm | 8717.22 | On request | RIB140-34
| S3546IL | Smart 3G-D3, Isobutylene VOC PID 0-200 ppm | 8717.22 | On request | RIB141-34
| S5123IL | Smart 3G-D3, Isobutylene VOC PID 0-2000 ppm | 8717.22 | On request | RIB142-34
| S5022BZ | Smart 3G-D3, Benzene VOC PID 0-10 ppm | 8717.22 | On request | RIB140-34
| S5026BZ | Smart 3G-D3, Benzene VOC PID 0-100 ppm | 8717.22 | On request | RIB141-34
| S5030BZ | Smart 3G-D3, Benzene VOC PID 0-1000 ppm | 8717.22 | On request | RIB142-34
| S5022TO | Smart 3G-D3, Toluene VOC PID 0-10 ppm | 8717.22 | On request | RIB140-34
| S5026TO | Smart 3G-D3, Toluene VOC PID 0-100 ppm | 8717.22 | On request | RIB141-34
| S5030TO | Smart 3G-D3, Toluene VOC PID 0-1000 ppm | 8717.22 | On request | RIB142-34
Accessories
| Product Code | Description | PVP
| STS/CSW+ | Calibration software for SMART3G, SMART3NC & SMART3R detectors consisting of a USB key, a USB/RS485 converter and the STS/IDI card. The software runs on Windows XP, Windows 7 e Windows 10. The calibration software allows to modify the alarm thresholds and the measuring range, and allows adjusting the Zero, Span and 4-20 mA values via a user-friendly menu, either in situ or in laboratory. | 620.46
| STS/IDI | Plug-in interface for RS485 serial communication suitable for SMART3G | 84.24
| STGD/AD3 | Adapter for all sensor heads (exception made for S2096__ & S2097__) to upgrade the protection from G to GD | 103.68

## SMART 3G-C3
Auto+Hide+Values
SMART 3G-C3 For Zone 2 category 3 areas
The SMART 3G-C3 gas detectors are designed to meet industrial requirements, enabling the monitoring of toxic, flammable, and refrigerant gases in areas classified as Zone 2, Category 3. Ideal for light industrial environments, the SMART 3G-C3 features one-man calibration using the STS/CKD+ calibration keypad. These detectors can reliably monitor flammable compounds (% LFL), toxic and refrigerant gases in ppm, as well as oxygen and CO2 levels (% by volume). Features: • Standard 4-20 mA 3 wire output • Optional 1 or 3 relay output • Optional RS485 interface for Modbus communication. • Ongoing system self diagnosis • ATEX: II 3G Ex db ec nC IIC T5 Gc • SIL 2 HW & SIL3 SW certified • IP55
Flammable Catalytic Sensors
| Product Code | Description | PVP | Sensor Head | Gas Cylinder
| S2096ME | Smart 3G-C3, Methane Cat 0-100% LEL | 822.96 | Kx097ME | SIB09
| S2096GP | Smart 3G-C3, LPG Cat 0-100% LEL | 822.96 | Kx097GP | SIB06
| S2096PR | Smart 3G-C3, Propane Cat 0-100% LEL | 822.96 | Kx097PR | SIB09
| S2096BU | Smart 3G-C3, Butane Cat 0-100% LEL | 822.96 | Kx097BU | SIB06
| S2096VB | Smart 3G-C3, Petrol Vapours Cat 0-100% LEL | 822.96 | Kx097VB | SIB06
| S2096H2 | Smart 3G-C3, Hydrogen Cat 0-100% LEL | 822.96 | Kx097H2 | SIB19
Flammable Infra-Red (IR) Sensors
| Product Code | Description | PVP | Sensor Head | Gas Cylinder
| S2623ME | Smart 3G-C3, Methane IR 0-100% LEL | 1675.08 | Kx624ME | SIB09
| S2653PR | Smart 3G-C3, Propane IR 0-100% LEL | 1675.08 | Kx654PR | SIB28
| S2653BU | Smart 3G-C3, Butane IR 0-100% LEL | 1675.08 | Kx654BU | SIB28
| S2653GP | Smart 3G-C3, LPG IR 0-100% LEL | 1675.08 | Kx654GP | SIB28
Toxic Electrochemical Sensors and Oxygen
| Product Code | Description | PVP | Sensor Head | Gas Cylinder
| S2640O2 | Smart 3G-C3, Oxygen O2 EC 0-25% Vol | 1150.2 | Kx641O2 | SIB62
| S2128O2 | Smart 3G-C3, Oxygen O2 EC 0-30% Vol | 1150.2 | Kx129O2 | SIB62
| S2130CO | Smart 3G-C3, Carbon Monoxide CO EC 0-500 ppm | 1150.2 | Kx131CO | SIB56
| S3224AM | Smart 3G-C3, Ammonia NH3 EC 0-100 ppm | 1673.46 | Kx225AM | RIB73-34
| S2132AM | Smart 3G-C3, Ammonia NH3 EC 0-200 ppm | 1673.46 | Kx133AM | RIB74-34
| S2134AM | Smart 3G-C3, Ammonia NH3 EC 0-1000 ppm | 1673.46 | Kx135AM | RIB76-34
| S3228AM | Smart 3G-C3, Ammonia NH3 EC 0-5000 ppm | 1673.46 | K3229AM | RIB70-34
| S2136HS | Smart 3G-C3, Hydrogen Sulphide H2S EC 0-50 ppm | 1150.2 | Kx137HS | RIB81-34
| S2144NO | Smart 3G-C3, Nitric Oxide NO EC 0-100 ppm | 1840.32 | Kx145NO | RIB81-34
| S3592H2 | Smart 3G-C3, Hydrogen H2 EC 0-2000 ppm | 1840.32 | K3593H2 | RIB92-34
| S4546H2 | Smart 3G-C3, Hydrogen H2 EC 0-3000 ppm | 1840.32 | K4547H2 | RIB92-34
CO2 Infra-Red (IR) Sensors
| Product Code | Description | PVP | Sensor Head | Gas Cylinder
| S2444CO2 | Smart 3G-C3, Carbon Dioxide CO2 IR 0-2% Vol | 1675.08 | Kx444CO2 | RIB110-34
| S2649CO2 | Smart 3G-C3, Carbon Dioxide CO2 IR 0-5% Vol | 1675.08 | Kx651CO2 | SIB42
| S2650CO2 | Smart 3G-C3, Carbon Dioxide CO2 IR 0-30% Vol | 1675.08 | Kx652CO2 | SIB44
| S2443CO2 | Smart 3G-C3, Carbon Dioxide CO2 IR 0-5000 ppm | 1675.08 | Kx445CO2 | SIB48
| S2629CO2 | Smart 3G-C3, Carbon Dioxide CO2 IR 0-100% Vol | 1675.08 | Kx630CO2 | SIB48
Infra-Red (IR) Sensors - Refrigerant Gases A1, A2L and SF6
| Product Code | Description | PVP | Sensor Head | Gas Cylinder
| S4432R454B | Smart 3G-C3, R454b IR 0-100% LEL | 2888.46 | K4433R454B | SIB507
| S4826R454C | Smart 3G-C3, R454c IR 0-100% LEL | 2888.46 | K4827R454C | On Request
| S3424R32 | Smart 3G-C3, R32 IR 0-100% LEL | 2888.46 | K3425R32 | SIB512
| S3428R1234YF | Smart 3G-C3, R1234yf IR 0-100% LEL | 2888.46 | K3429R1234YF | SIB501
| S3432R1234ZE | Smart 3G-C3, R1234ze IR 0-100% LEL | 2888.46 | K3433R1234ZE | SIB502
| S3942R32 | Smart 3G-C3, R32 IR 0-2000 ppm | 2888.46 | Kx944R32 | RIB148-34
| S3595R134A | Smart 3G-C1, R134a IR 0-2000 ppm | 2888.46 | K3596R134A | RIB144-34
| S3598SF6 | Smart 3G-C6, SF6 IR 0-2000 ppm | 2888.46 | K3599SF6 | RIB149-34
| S3020R125 | Smart 3G-C1, R125 IR 0-2000 ppm | 2888.46 | K3656R125 | RIB130-34
| S3021R1234YF | Smart 3G-C1, R1234yf IR 0-2000 ppm | 2888.46 | K3657R1234YF | RIB145-34
| S3025R1234ZE | Smart 3G-C1, R1234ze IR 0-2000 ppm | 2888.46 | K3948R1234ZE | RIB143-34
| S3022R404A | Smart 3G-C4, R404a IR 0-2000 ppm | 2888.46 | K3658R404A | RIB131-34
| S3023R407A | Smart 3G-C4, R407a IR 0-2000 ppm | 2888.46 | K3659R407A | RIB132-34
| S3918R407F | Smart 3G-C4, R407f IR 0-2000 ppm | 2888.46 | K3920R407F | RIB133-34
| S3921R449A | Smart 3G-C4, R449 IR 0-2000 ppm | 2888.46 | K3923R449A | RIB134-34
| S3024R507A | Smart 3G-C5, R507a IR 0-2000 ppm | 2888.46 | K3660R507A | RIB135-34
| S3016R410A | Smart 3G-C4, R410a IR 0-2000 ppm | 2888.46 | K3017R410A | RIB147-34
| S3950R417A | Smart 3G-C4, R417a IR 0-2000 ppm | 2888.46 | K3952R417A | RIB136-34
| S3954R448 | Smart 3G-C4, R448 IR 0-2000 ppm | 2888.46 | K3956R448 | RIB137-34
VOC Photoionisation (PID) Sensors
| Product Code | Description | PVP | Sensor Head | Gas Cylinder
| S3540IL | Smart 3G-C3, Isobutylene VOC PID 0-20 ppm | 8179.38 | On request | RIB140-34
| S3544IL | Smart 3G-C3, Isobutylene VOC PID 0-200 ppm | 8179.38 | On request | RIB141-34
| S5121IL | Smart 3G-C3, Isobutylene VOC PID 0-2000 ppm | 8179.38 | On request | RIB142-34
| S5020BZ | Smart 3G-C3, Benzene VOC PID 0-10 ppm | 8179.38 | On request | RIB140-34
| S5024BZ | Smart 3G-C3, Benzene VOC PID 0-100 ppm | 8179.38 | On request | RIB141-34
| S5028BZ | Smart 3G-C3, Benzene VOC PID 0-1000 ppm | 8179.38 | On request | RIB142-34
| S5020TO | Smart 3G-C3, Toluene VOC PID 0-10 ppm | 8179.38 | On request | RIB140-34
| S5024TO | Smart 3G-C3, Toluene VOC PID 0-100 ppm | 8179.38 | On request | RIB141-34
| S5028TO | Smart 3G-C3, Toluene VOC PID 0-1000 ppm | 8179.38 | On request | RIB142-34
Accessories
| Product Code | Description | PVP
| STS/CSW+ | Calibration software for SMART3G, SMART3NC & SMART3R detectors consisting of a USB key, a USB/RS485 converter and the STS/IDI card. The software runs on Windows XP, Windows 7 e Windows 10. The calibration software allows to modify the alarm thresholds and the measuring range, and allows adjusting the Zero, Span and 4-20 mA values via a user-friendly menu, either in situ or in laboratory. | 626.94
| STS/IDI | Plug-in RS485 interface, to make SMART3G series gas detectors communicate in Modbus protocol with MULTISCAN++S1 & S2, MULTISCAN 8+ & STPL4+D control panels. | 84.24
| STS/CKD+ | Enhanced calibration keypad with 7-digit display and SD card slot. Compatible with Smart series for operations like zero/span calibration and monitor 4-20 mA signals. | 515.16
| STS/CKD-OLED | Enhanced calibration keypad with OLED display and SD-CARD slot. It can be connected to all SMART detectors to adjust the Zero, Span and 4-20 mA values. This version offers enhanced features like: - visualization of the instantaneous concentration in graphic format; - possibility to monitor the detector trend and download the data stored in the SD-CARD - detector’s firmware update via SD-CARD (feature not enabled on old SMART3G and NC detector) | 780.84
| STS1REL | 1-Relay card for SMART3G. Relay at 12÷24V | 84.24
| STS3REL | 3-Relay card for SMART3G. Relay at 12÷24V | 124.74
| STGD/AD2 | Adapter for S2096__ & S2097__ sensor heads to upgrade the protection from G to GD | 103.68
| STGD/AD3 | Adapter for all sensor heads (exception made for S2096__ & S2097__) to upgrade the protection from G to GD | 103.68

## SMART 3G-GRI
Auto+Hide+Values
SMART 3G-GR1 Gas Detectors for Group I Mines and Tunnels
The SMART 3G-GR1 gas detectors are designed to detect flammable and toxic gases in mines, tunnels, and areas classified as Group I. The SMART 3GD-GR1 variant, which includes a front display, features a 4-digit back-lit display for gas concentration readings, five mode status LEDs, and a high-visibility multi-color LED light ring. This version also offers non-intrusive magnetic calibration for precise and easy adjustments via Hall-effect switches, eliminating the need to open the instrument and declassify the area. Both versions of the SMART 3G-GR1 detectors are ATEX certified for Group I. Features: • Stainless steel enclosure • Non-intrusive one man calibration via Hall effect switches (display version) • 4-digit display and multi-color LED light ring (display version) • Standard 4-20 mA 3 wire output • Optional 3 relay output (standard 3-relay output in display version) • Optional RS485 interface for Modbus communication. • Zero tracking to maintain the zero point steady from possible drifts • Ongoing system self diagnosis • ATEX I M2 Ex d I Mb • SIL 2HW & SIL3 SW certified
With Display
| Product Code | Description | PVP | Sensor Head | Gas Cylinder
| S3515ME | Smart 3G-GR1, Display, Methane Cat. 0-100% LEL, 4-20 mA | 2656.8 | K3914ME | SIB09
| S3912ME | Smart 3G-GR1, Display, Methane IR 0-100% LEL, 4-20 mA | 3173.58 | K3916ME | SIB09
| S3911O2 | Smart 3G-GR1, Display, Oxygen EC 0-25% Vol., 4-20 mA | 2998.62 | K3915O2 | SIB62
| S3514CO | Smart 3G-GR1, Display, Carbon Monoxide EC 0-500 ppm, 4-20 mA | 2998.62 | K3913CO | SIB56
No display
| Product Code | Description | PVP | Sensor Head | Gas Cylinder
| S3914ME | Smart 3G-GR1, No Display, Methane Cat. 0-100% LEL, 4-20 mA | 2046.06 | K3914ME | SIB09
| S3916ME | Smart 3G-GR1, No Display, Methane IR 0-100% LEL, 4-20 mA | 2896.56 | K3916ME | SIB09
| S3915O2 | Smart 3G-GR1, No Display, Oxygen EC 0-25% Vol., 4-20 mA | 2357.1 | K3915O2 | SIB62
| S3913CO | Smart 3G-GR1, No Display, Carbon Monoxide EC 0-500 ppm, 4-20 mA | 2357.1 | K3913CO | SIB56

## COVER SMART S
#VALUE!

## SMART S-SS
SMART S-SS
The SMART S-SS line is designed to meet the most demanding industrial requirements, enabling the monitoring of flammable, toxic, and refrigerant gases in harsh environments. Features: - Electronic with1 output RS+4-20 mA, - LCD or LED display - Optional 3 relay output - Optional Modem Hart - Aluminum or Stainless Steel Enclosure - SIL2 Hardware and SIL3 Software Certified - ATEX certified for zone 1 as II2G Ex db IIC T5 Gb
Product Configurator
Transmitter | SSS3A11 | Smart S-SS Transmitter, 3A Alum Enclosure, RS-485, 4-20 mA, LCD Disp | 1328.4
Sensor Head | S4003ME | Smart S Sensor Mod, Methane Cat 0-100% LEL | 928.26
Use the configurator to generate part number, price and description.
Configured Code No. |  | PVP
SSS3A11-S4003ME |  | 3656.34
0.6
Smart S-SS Transmitters 3A Aluminum Enclosure complete with electronics
Product Code | Description | PVP
SSS3A11 | Smart S-SS Transmitter, 3A Alum Enclosure, RS-485, 4-20 mA, LCD Disp | 1328.4
SSS3A12 | Smart S-SS Transmitter, 3A Alum Enclosure, RS-485, 4-20 mA, LED Disp | 1644.3
SSS3A13 | Smart S-SS Transmitter, 3A Alum Enclosure, RS-485, 4-20 mA, LCD Disp, Relay Card | 1438.56
SSS3A14 | Smart S-SS Transmitter, 3A Alum Enclosure, RS-485, 4-20 mA, LED Disp, Relay Card | 1739.88
SSS3A17 | Smart S-SS Transmitter, 3A Alum Enclosure, RS-485, 4-20 mA, LCD Disp, Relay Card, HART | 2366.82
SSS3A18 | Smart S-SS Transmitter, 3A Alum Enclosure, RS-485, 4-20 mA, LED Disp, Relay Card, HART | 2697.3
Smart S-SS Transmitters 4A Aluminum Enclosure complete with electronics
Product Code | Description | PVP
SSS4A11 | Smart S-SS Transmitter, 4A Alum Enclosure, RS-485, 4-20 mA, LCD Disp | 1438.56
SSS4A12 | Smart S-SS Transmitter, 4A Alum Enclosure, RS-485, 4-20 mA, LED Disp | 1739.88
SSS4A13 | Smart S-SS Transmitter, 4A Alum Enclosure, RS-485, 4-20 mA, LCD Disp, Relay Card | 1547.1
SSS4A14 | Smart S-SS Transmitter, 4A Alum Enclosure, RS-485, 4-20 mA, LED Disp, Relay Card | 1848.42
SSS4A17 | Smart S-SS Transmitter, 4A Alum Enclosure, RS-485, 4-20 mA, LCD Disp, Relay Card, HART | 2475.36
SSS4A18 | Smart S-SS Transmitter, 4A Alum Enclosure, RS-485, 4-20 mA, LED Disp, Relay Card, HART | 2813.94
Smart S-SS Transmitters 2S Stainless Steel Enclosure complete with electronics
Product Code | Description | PVP
SSS2S11 | Smart S-SS Transmitter, 2S SS Enclosure, RS-485, 4-20 mA, LCD Disp | 2055.78
SSS2S12 | Smart S-SS Transmitter, 2S SS Enclosure, RS-485, 4-20 mA, LED Disp | 2366.82
SSS2S13 | Smart S-SS Transmitter, 2S SS Enclosure, RS-485, 4-20 mA, LCD Disp, Relay Card | 2151.36
SSS2S14 | Smart S-SS Transmitter, 2S SS Enclosure, RS-485, 4-20 mA, LED Disp, Relay Card | 2465.64
SSS2S17 | Smart S-SS Transmitter, 2S SS Enclosure, RS-485, 4-20 mA, LCD Disp, Relay Card, HART | 3149.28
SSS2S18 | Smart S-SS Transmitter, 2S SS Enclosure, RS-485, 4-20 mA, LED Disp, Relay Card, HART | 3479.76
Smart S-SS Sensor Modules
Flammable Catalytic Sensors
Product Code | Description | PVP | Gas Cylinder
S4003ME | Smart S Sensor Mod, Methane Cat 0-100% LEL | 928.26 | SIB09
S4001ME | Smart S Sensor Mod, Methane Cat 0-100% LEL (5% Vol) | 1438.56 | SIB09
S4003GP | Smart S Sensor Mod, LPG Cat 0-100% LEL | 928.26 | SIB06
S4003PR | Smart S Sensor Mod, Propane Cat 0-100% LEL | 928.26 | SIB06
S4003BU | Smart S Sensor Mod, Butane Cat 0-100% LEL | 928.26 | SIB06
S4003VB | Smart S Senor Mod, Petrol Vapours Cat 0-100% LEL | 928.26 | SIB06
S4003AT | Smart S Sensor Mod, Acetone Cat 0-100% LEL | 928.26 | SIB021
S4037AC | Smart S Sensor Mod, Acetylene Cat 0-100% LEL | 928.26 | SIB19
S4003AM | Smart S Sensor Mod, Ammonia Cat 0-100% LEL | 928.26 | SIB06
S4003CP | Smart S Sensor Mod, Cyclopentane Cat 0-100% LEL | 928.26 | SIB06
S4003EN | Smart S Sensor Mod, Ethane Cat 0-100% LEL | 928.26 | SIB06
S4003AE | Smart S Senor Mod, Ethyl Acetate Cat 0-100% LEL | 928.26 | SIB06
S4003ET | Smart S Senor Mod, Ethyl Alcohol Cat 0-100% LEL | 928.26 | SIB09
S4003EL | Smart S Sensor Mod, Ethylene Cat 0-100% LEL | 928.26 | SIB06
S4003EP | Smart S Sensor Mod, Heptane Cat 0-100% LEL | 928.26 | SIB06
S4003ES | Smart S Sensor Mod, Hexane Cat 0-100% LEL | 928.26 | SIB06
S4003H2 | Smart S Sensor Mod, Pentane Cat 0-100% LEL | 928.26 | SIB19
S4003IB | Smart S Senor Mod, Iso Butane Cat 0-100% LEL | 928.26 | SIB06
S4002P8 | Smart S Sensor Mod, JP8 Cat 0-100% LEL | 928.26 | SIB36
S4003MT | Smart S Senor Mod, Methyl Alcohol Cat 0-100% LEL | 928.26 | SIB09
S4003MK | Smart K Senor Mod, Methyl Ethyl Cat 0-100% LEL | 928.26 | SIB06
S4003PE | Smart S Sensor Mod, Pentane Cat 0-100% LEL | 928.26 | SIB06
S4003PP | Smart S Sensor Mod, Propylene Cat 0-100% LEL | 928.26 | SIB09
S4003PN | Smart S Senor Mod, Propyl Alcohol Cat 0-100% LEL | 928.26 | SIB021
S4003ST | Smart S Sensor Mod, Styrene Cat 0-100% LEL | 928.26 | SIB021
S4003TO | Smart S Sensor Mod, Toluene Cat 0-100% LEL | 928.26 | SIB06
S4003XI | Smart S Sensor Mod, Xylene Cat 0-100% LEL | 928.26 | SIB06
Flammable Infra-Red (IR) Sensors
Product Code | Description | PVP | Gas Cylinder
S4019ME | Smart S Sensor Mod, Methane Cat 0-100% LEL | 1438.56 | SIB09
S4004PR | Smart S Sensor Mod, Propane IR 0-100% LEL | 1438.56 | SIB28
S4004BU | Smart S Sensor Mod, Butane IR 0-100% LEL | 1438.56 | SIB28
S4004AT | Smart S Sensor Mod, Acetone IR 0-100% LEL | 1438.56 | SIB22
S4004BN | Smart S Sensor Mod, Butadiene IR 0-100% LEL | 1438.56 | SIB26
S4004CE | Smart S Sensor Mod, Cyclo Hexane IR 0-100% LEL | 1438.56 | SIB26
S4004DT | Smart S Sensor Mod, Dimethyl Ether IR 0-100% LEL | 1438.56 | SIB28
S4004EN | Smart S Sensor Mod, Ethane IR 0-100% LEL | 1438.56 | SIB28
S4004AE | Smart S Sensor Mod, Ethyl Acetate IR 0-100% LEL | 1438.56 | SIB26
S4004ET | Smart S Sensor Mod, Ethyl Alcohol IR 0-100% LEL | 1438.56 | SIB28
S4004EP | Smart S Sensor Mod, Heptane IR 0-100% LEL | 1438.56 | SIB26
S4004ES | Smart S Sensor Mod, Hexane IR 0-100% LEL | 1438.56 | SIB26
S4004IB | Smart S Sensor Mod, Iso Butane IR 0-100% LEL | 1438.56 | SIB26
S4004IL | Smart S Sensor Mod, Isobutylene IR 0-100% LEL | 1438.56 | SIB26
S4004IP | Smart S Sensor Mod, Iso Propyl Alcohol IR 0-100% LEL | 1438.56 | SIB26
S4004PE | Smart S Sensor Mod, Pentane IR 0-100% LEL | 1438.56 | SIB28
S4004PP | Smart S Sensor Mod, Propylene IR 0-100% LEL | 1438.56 | SIB26
S4004TO | Smart S Sensor Mod, Toluene IR 0-100% LEL | 1438.56 | SIB22
SF6 & Refrigerant Infra-Red (IR) Sensors
Product Code | Description | PVP | Gas Cylinder
S4004DE | Smart S Sensor Mod, R152a IR 0-100% LEL | 2587.14 | SIB26
S4157R454B | Smart S Sensor Mod, R454b IR 0-100% LEL | 2587.14 | SIB507
S4065R32 | Smart S Sensor Mod, R32 IR 0-100% LEL | 2587.14 | SIB512
S4066R1234YF | Smart S Sensor Mod, R1234yf IR 0-100% LEL | 2587.14 | SIB501
S4067R1234ZE | Smart S Sensor Mod, R1234ze IR 0-100% LEL | 2587.14 | SIB502
S4152R32 | Smart S Sensor Mod, R32 IR 0-2000 ppm | 2587.14 | RIB148-34
S4026R134A | Smart S Sensor Mod, R134a IR 0-2000 ppm | 2587.14 | RIB144-34
S4027SF6 | Smart S Sensor Mod, SF6 IR 0-2000 ppm | 2587.14 | RIB149-34
S4031R125 | Smart S Sensor Mod, R125 IR 0-2000 ppm | 2587.14 | RIB130-34
S4032R1234YF | Smart S Sensor Mod, R1234yf IR 0-2000 ppm | 2587.14 | RIB145-34
S4153R1234ZE | Smart S Sensor Mod, R1234ze IR 0-2000 ppm | 2587.14 | RIB143-34
S4033R404A | Smart S Sensor Mod, R404a IR 0-2000 ppm | 2587.14 | RIB131-34
S4034R407A | Smart S Sensor Mod, R407a IR 0-2000 ppm | 2587.14 | RIB132-34
S4034R407F | Smart S Sensor Mod, R407f IR 0-2000 ppm | 2587.14 | RIB133-34
S4154R449A | Smart S Sensor Mod, R449a IR 0-2000 ppm | 2587.14 | RIB134-34
S4035R507A | Smart S Sensor Mod, R507a IR 0-2000 ppm | 2587.14 | RIB135-34
S4036R410A | Smart S Sensor Mod, R410a IR 0-2000 ppm | 2587.14 | RIB147-34
S4155R417A | Smart S Sensor Mod, R417a IR 0-2000 ppm | 2587.14 | RIB136-34
S4156R448 | Smart S Sensor Mod, R448 IR 0-2000 ppm | 2587.14 | RIB137-34
CO2 Infra-Red (IR) Sensors
Product Code | Description | PVP | Gas Cylinder
S4015CO2 | Testa sensore ANIDRIDE CARBONICA (CO2), 0-2% Vol. | 1438.56 | SIB42
S4016CO2 | Testa sensore ANIDRIDE CARBONICA (CO2), 0-5% Vol. | 1438.56 | SIB44
S4017CO2 | Testa sensore ANIDRIDE CARBONICA (CO2), 0-30% Vol. | 1438.56 | SIB48
Toxic Electrochemical Sensors
Product Code | Description | PVP | Gas Cylinder
S4005O2 | Smart S Sensor Mod, Oxygen O2 EC 0-25% Vol | 1135.62 | SIB62
S4166CO | Smart S Sensor Mod, Carbon Monoxide CO EC 0-100 ppm | 1135.62 | SIB50
S4006CO | Smart S Sensor Mod, Carbon Monoxide CO EC 0-500 ppm | 1135.62 | SIB56
S4021CO | Smart S Sensor Mod, Carbon Monoxide CO EC 0-1000 ppm | 1135.62 | SIB58
S4022AM | Smart S Sensor Mod, Ammonia NH3 EC 0-100 ppm | 1438.56 | RIB73-34
S4007AM | Smart S Sensor Mod, Ammonia NH3 EC 0-200 ppm | 1438.56 | RIB74-34
S4008AM | Smart S Sensor Mod, Ammonia NH3 EC 0-1000 ppm | 1438.56 | RIB76-34
S4009HS | Smart S Sensor Mod, Hydrogen Sulphide H2S EC 0-50 ppm | 1135.62 | RIB81-34
S4028HS | Smart S Sensor Mod, Hydrogen Sulphide H2S EC 0-20 ppm | 1135.62 | RIB82-34
S4029HS | Smart S Sensor Mod, Hydrogen Sulphide H2S EC 0-100 ppm | 1135.62 | RIB81-34
S4010NO | Smart S Sensor Mod, Nitric Oxide NO EC 0-100 ppm | 1438.56 | RIB117-34
S4089H2 | Smart S Sensor Mod, Hydrogen H2 EC 0-1000 ppm | 1438.56 | RIB92-34
S4168H2 | Smart S Sensor Mod, Hydrogen H2 EC 0-2000 ppm | 1438.56 | RIB92-34
S4169H2 | Smart S Sensor Mod, Hydrogen H2 EC 0-3000 ppm | 1438.56 | RIB92-34
VOC Photoionisation (PID) Sensors
Product Code | Description | PVP | Gas Cylinder
S4167IL | Smart S Sensor Mod, Isobutylene PID 0-20 ppm | On request | RIB140-34
S4158IL | Smart S Sensor Mod, Isobutylene PID 0-200 ppm | On request | RIB141-34
S4159IL | Smart S Sensor Mod, Isobutylene PID 0-2000 ppm | On request | RIB142-34
S4075BZ | Smart S Sensor Mod, Benzene PID 0-10 ppm | On request | RIB140-34
S4160BZ | Smart S Sensor Mod, Benzene PID 0-100 ppm | On request | RIB141-34
S4161BZ | Smart S Sensor Mod, Benzene PID 0-1000 ppm | On request | RIB142-34
S4075TO | Smart S Sensor Mod, Toluene PID 0-10 ppm | On request | RIB140-34
S4160TO | Smart S Sensor Mod, Toluene PID 0-100 ppm | On request | RIB141-34
S4161TO | Smart S Sensor Mod, Toluene PID 0-1000 ppm | On request | RIB142-34
Accessories
Product Code | Description | PVP
STS/CKD+ | Enhanced calibration keypad with 7-digit display and SD card slot. Compatible with Smart series for operations like zero/span calibration and monitor 4-20 mA signals. | 515.16
STS/CKD-OLED | Enhanced calibration keypad with OLED display and SD-CARD slot. Can be connected to all SMART detectors to adjust Zero, Span and 4-20 mA values. Features: visualization of instantaneous concentration in graphic format; Monitor detector trend & download data stored in the SD-CARD - Detector firnware can be updated via SD-card (feature not enabled on old SMART3G & NC) | 780.84

## SMART S-MS
Auto+Hide+Values
SMART S-MS
The SMART S-MS line is engineered to meet the most stringent industrial requirements, enabling the monitoring of flammable, toxic, and refrigerant gases in harsh environments. The detector is available in either alloy or stainless steel housing for enhanced durability. Features: - Electronic with1 output RS+4-20 mA, - Optional 3 relay output - Optional Modem Hart - Aluminum or Stainless Steel Enclosure - SIL2 Hardware and SIL3 Software Certified - ATEX certified for zone 1 as II2G Ex db IIC T5 Gb - IP65
Product Configurator
| Transmitter | SMS3A11 | Smart S-MS Transmitter, 3A Alum Encl, RS-485, 4-20 mA | 928.26
| Sensor | S4032R1234YF | Smart S Sensor Mod, R1234yf IR 0-2000 ppm | 2587.14
Use the configurator to generate part number, price and description.
| Cofnigured Part No. |  | PVP
| SMS3A11-S4032R1234YF |  | 5694.3
Smart S-MS Transmitter, 3A Alum Encl, RS-485, 4-20 mA - Smart S Sensor Mod, R1234yf IR 0-2000 ppm
Smart S-MS Transmitters
Smart S-MS Transmitters 3A Aluminum Enclosure complete with electronics
| Product Code | Description | PVP
| SMS3A11 | Smart S-MS Transmitter, 3A Alum Encl, RS-485, 4-20 mA | 928.26
| SMS3A13 | Smart S-MS Transmitter, 3A Alum Encl, RS-485, 4-20 mA, Relay Card | 1028.7
| SMS3A17 | Smart S-MS Transmitter, 3A Alum Encl, RS-485, 4-20 mA, Relay Card, HART | 1739.88
Smart S-MS Transmitters 4A Aluminum Enclosure complete with electronics
| Product Code | Description | PVP
| SMS4A11 | Smart S-MS Transmitter, 4A Alum Encl, RS-485, 4-20 mA | 1028.7
| SMS4A13 | Smart S-MS Transmitter, 4A Alum Encl, RS-485, 4-20 mA, Relay Card | 1135.62
| SMS4A17 | Smart S-MS Transmitter, 4A Alum Encl, RS-485, 4-20 mA, Relay Card, HART | 2033.1
Smart S-MS Transmitters 2S Stainless Steel Enclosure complete with electronics
| Product Code | Description | PVP
| SMS2S11 | Smart S-MS Transmitter, 2S SS Encl, RS-485, 4-20 mA | 1644.3
| SMS2S13 | Smart S-MS Transmitter, 2S SS Encl, RS-485, 4-20 mA, Relay Card | 1793.34
| SMS2S17 | Smart S-MS Transmitter, 2S SS Encl, RS-485, 4-20 mA, Relay Card, HART | 2697.3
Smart S-MS Sensor Modules
Flammable Catalytic Sensors
| Product Code | Description | PVP | Gas Cylinder
| S4003ME | Smart S Sensor Mod, Methane Cat 0-100% LEL | 928.26 | SIB09
| S4001ME | Smart S Sensor Mod, Methane Cat 0-100% LEL (5% Vol.) | 1438.56 | SIB09
| S4003GP | Smart S Sensor Mod, LPG Cat 0-100% LEL | 928.26 | SIB06
| S4003PR | Smart S Sensor Mod, Propane Cat 0-100% LEL | 928.26 | SIB06
| S4003BU | Smart S Sensor Mod, Butane Cat 0-100% LEL | 928.26 | SIB06
| S4003VB | Smart S Senor Mod, Petrol Vapours Cat 0-100% LEL | 928.26 | SIB06
| S4003AT | Smart S Sensor Mod, Acetone Cat 0-100% LEL | 928.26 | SIB021
| S4037AC | Smart S Sensor Mod, Acetylene Cat 0-100% LEL | 928.26 | SIB19
| S4003AM | Smart S Sensor Mod, Ammonia Cat 0-100% LEL | 928.26 | SIB06
| S4003CP | Smart S Sensor Mod, Cyclopentane Cat 0-100% LEL | 928.26 | SIB06
| S4003EN | Smart S Sensor Mod, Ethane Cat 0-100% LEL | 928.26 | SIB06
| S4003AE | Smart S Senor Mod, Ethyl Acetate Cat 0-100% LEL | 928.26 | SIB06
| S4003ET | Smart S Senor Mod, Ethyl Alcohol Cat 0-100% LEL | 928.26 | SIB09
| S4003EL | Smart S Sensor Mod, Ethylene Cat 0-100% LEL | 928.26 | SIB06
| S4003EP | Smart S Sensor Mod, Heptane Cat 0-100% LEL | 928.26 | SIB06
| S4003ES | Smart S Sensor Mod, Hexane Cat 0-100% LEL | 928.26 | SIB06
| S4003H2 | Smart S Sensor Mod, Hydrogen Cat 0-100% LEL | 928.26 | SIB19
| S4003IB | Smart S Senor Mod, Iso Butane Cat 0-100% LEL | 928.26 | SIB06
| S4002P8 | Smart S Sensor Mod, JP8 Cat 0-100% LEL | 928.26 | SIB36
| S4003MT | Smart S Senor Mod, Methyl Alcohol Cat 0-100% LEL | 928.26 | SIB09
| S4003MK | Smart K Senor Mod, Methyl Ethyl Cat 0-100% LEL | 928.26 | SIB06
| S4003PE | Smart S Sensor Mod, Pentane Cat 0-100% LEL | 928.26 | SIB06
| S4003PP | Smart S Sensor Mod, Propylene Cat 0-100% LEL | 928.26 | SIB09
| S4003PN | Smart S Senor Mod, Propyl Alcohol Cat 0-100% LEL | 928.26 | SIB021
| S4003ST | Smart S Sensor Mod, Styrene Cat 0-100% LEL | 928.26 | SIB021
| S4003TO | Smart S Sensor Mod, Toluene Cat 0-100% LEL | 928.26 | SIB06
| S4003XI | Smart S Sensor Mod, Xylene Cat 0-100% LEL | 928.26 | SIB06
Flammable Infra-Red (IR) Sensors
| Product Code | Description | PVP | Gas Cylinder
| S4019ME | Smart S Sensor Mod, Methane IR 0-100% LEL (4.4% Vol) | 1438.56 | SIB09
| S4004PR | Smart S Sensor Mod, Propane IR 0-100% LEL | 1438.56 | SIB28
| S4004BU | Smart S Sensor Mod, Butane IR 0-100% LEL | 1438.56 | SIB28
| S4004AT | Smart S Sensor Mod, Acetone IR 0-100% LEL | 1438.56 | SIB22
| S4004BN | Smart S Sensor Mod, Butadiene IR 0-100% LEL | 1438.56 | SIB26
| S4004CE | Smart S Sensor Mod, Cyclo Hexane IR 0-100% LEL | 1438.56 | SIB26
| S4004DT | Smart S Sensor Mod, Dimethyl Ether IR 0-100% LEL | 1438.56 | SIB28
| S4004EN | Smart S Sensor Mod, Ethane IR 0-100% LEL | 1438.56 | SIB28
| S4004AE | Smart S Sensor Mod, Ethyl Acetate IR 0-100% LEL | 1438.56 | SIB26
| S4004ET | Smart S Sensor Mod, Ethyl Alcohol IR 0-100% LEL | 1438.56 | SIB28
| S4004EP | Smart S Sensor Mod, Heptane IR 0-100% LEL | 1438.56 | SIB26
| S4004ES | Smart S Sensor Mod, Hexane IR 0-100% LEL | 1438.56 | SIB26
| S4004IB | Smart S Sensor Mod, Iso Butane IR 0-100% LEL | 1438.56 | SIB26
| S4004IL | Smart S Sensor Mod, Isobutylene IR 0-100% LEL | 1438.56 | SIB26
| S4004IP | Smart S Sensor Mod, Iso Propyl Alcohol IR 0-100% LEL | 1438.56 | SIB26
| S4004PE | Smart S Sensor Mod, Pentane IR 0-100% LEL | 1438.56 | SIB28
| S4004PP | Smart S Sensor Mod, Propylene IR 0-100% LEL | 1438.56 | SIB26
| S4004TO | Smart S Sensor Mod, Toluene IR 0-100% LEL | 1438.56 | SIB22
SF6 & Refrigerant Infra-Red (IR) Sensors
| Product Code | Description | PVP | Gas Cylinder
| S4004DE | Smart S Sensor Mod, R152a IR 0-100% LEL | 2587.14 | SIB26
| S4157R454B | Smart S Sensor Mod, R454b IR 0-100% LEL | 2587.14 | SIB507
| S4065R32 | Smart S Sensor Mod, R32 IR 0-100% LEL | 2587.14 | SIB512
| S4066R1234YF | Smart S Sensor Mod, R1234yf IR 0-100% LEL | 2587.14 | SIB501
| S4067R1234ZE | Smart S Sensor Mod, R1234ze IR 0-100% LEL | 2587.14 | SIB502
| S4152R32 | Smart S Sensor Mod, R32 IR 0-2000 ppm | 2587.14 | RIB148-34
| S4026R134A | Smart S Sensor Mod, R134a IR 0-2000 ppm | 2587.14 | RIB144-34
| S4027SF6 | Smart S Sensor Mod, SF6 IR 0-2000 ppm | 2587.14 | RIB149-34
| S4031R125 | Smart S Sensor Mod, R125 IR 0-2000 ppm | 2587.14 | RIB130-34
| S4032R1234YF | Smart S Sensor Mod, R1234yf IR 0-2000 ppm | 2587.14 | RIB145-34
| S4153R1234ZE | Smart S Sensor Mod, R1234ze IR 0-2000 ppm | 2587.14 | RIB143-34
| S4033R404A | Smart S Sensor Mod, R404a IR 0-2000 ppm | 2587.14 | RIB131-34
| S4034R407A | Smart S Sensor Mod, R407a IR 0-2000 ppm | 2587.14 | RIB132-34
| S4034R407F | Smart S Sensor Mod, R407f IR 0-2000 ppm | 2587.14 | RIB133-34
| S4154R449A | Smart S Sensor Mod, R449a IR 0-2000 ppm | 2587.14 | RIB134-34
| S4035R507A | Smart S Sensor Mod, R507a IR 0-2000 ppm | 2587.14 | RIB135-34
| S4036R410A | Smart S Sensor Mod, R410a IR 0-2000 ppm | 2587.14 | RIB147-34
| S4155R417A | Smart S Sensor Mod, R417a IR 0-2000 ppm | 2587.14 | RIB136-34
| S4156R448 | Smart S Sensor Mod, R448 IR 0-2000 ppm | 2587.14 | RIB137-34
CO2 Infra-Red (IR) Sensors
| Product Code | Description | PVP | Gas Cylinder
| S4015CO2 | Smart S Sensor Mod, Carbon Dioxide CO2 0-2% Vol | 1438.56 | SIB42
| S4016CO2 | Smart S Sensor Mod, Carbon Dioxide CO2 0-5% Vol | 1438.56 | SIB44
| S4017CO2 | Smart S Sensor Mod, Carbon Dioxide CO2 0-30% Vol | 1438.56 | SIB48
Toxic Electrochemical Sensors
| Product Code | Description | PVP | Gas Cylinder
| S4005O2 | Smart S Sensor Mod, Oxygen O2 EC 0-25% Vol | 1135.62 | SIB62
| S4166CO | Smart S Sensor Mod, Carbon Monoxide CO EC 0-100 ppm | 1135.62 | SIB50
| S4006CO | Smart S Sensor Mod, Carbon Monoxide CO EC 0-500 ppm | 1135.62 | SIB56
| S4021CO | Smart S Sensor Mod, Carbon Monoxide CO EC 0-1000 ppm | 1135.62 | SIB58
| S4022AM | Smart S Sensor Mod, Ammonia NH3 EC 0-100 ppm | 1438.56 | RIB73-34
| S4007AM | Smart S Sensor Mod, Ammonia NH3 EC 0-200 ppm | 1438.56 | RIB74-34
| S4008AM | Smart S Sensor Mod, Ammonia NH3 EC 0-1000 ppm | 1438.56 | RIB76-34
| S4009HS | Smart S Sensor Mod, Hydrogen Sulphide H2S EC 0-50 ppm | 1135.62 | RIB81-34
| S4028HS | Smart S Sensor Mod, Hydrogen Sulphide H2S EC 0-20 ppm | 1135.62 | RIB82-34
| S4029HS | Smart S Sensor Mod, Hydrogen Sulphide H2S EC 0-100 ppm | 1135.62 | RIB81-34
| S4010NO | Smart S Sensor Mod, Nitric Oxide NO EC 0-100 ppm | 1438.56 | RIB117-34
| S4089H2 | Smart S Sensor Mod, Hydrogen H2 EC 0-1000 ppm | 1438.56 | RIB92-34
| S4168H2 | Smart S Sensor Mod, Hydrogen H2 EC 0-2000 ppm | 1438.56 | RIB92-34
| S4169H2 | Smart S Sensor Mod, Hydrogen H2 EC 0-3000 ppm | 1438.56 | RIB92-34
VOC Photoionisation (PID) Sensors
| Product Code | Description | PVP | Gas Cylinder
| S4167IL | Smart S Sensor Mod, Isobutylene PID 0-20 ppm | 8143.74 | RIB140-34
| S4158IL | Smart S Sensor Mod, Isobutylene PID 0-200 ppm | 8143.74 | RIB141-34
| S4159IL | Smart S Sensor Mod, Isobutylene PID 0-2000 ppm | 8143.74 | RIB142-34
| S4075BZ | Smart S Sensor Mod, Benzene PID 0-10 ppm | 8143.74 | RIB140-34
| S4160BZ | Smart S Sensor Mod, Benzene PID 0-100 ppm | 8143.74 | RIB141-34
| S4161BZ | Smart S Sensor Mod, Benzene PID 0-1000 ppm | 8143.74 | RIB142-34
| S4075TO | Smart S Sensor Mod, Toluene PID 0-10 ppm | 8143.74 | RIB140-34
| S4160TO | Smart S Sensor Mod, Toluene PID 0-100 ppm | 8143.74 | RIB141-34
| S4161TO | Smart S Sensor Mod, Toluene PID 0-1000 ppm | 8143.74 | RIB142-34
Accessories
| Product Code | Description | PVP
| STS/CKD+ | Enhanced calibration keypad with 7-digit display and SD card slot. Compatible with Smart series for operations like zero/span calibration and monitor 4-20 mA signals. | 515.16
| STS/CKD-OLED | Enhanced calibration keypad with OLED display and SD-CARD slot. Can be connected to all SMART detectors to adjust Zero, Span and 4-20 mA values. Features: visualization of instantaneous concentration in graphic format; Monitor detector trend & download data stored in the SD-CARD - Detector firnware can be updated via SD-card (feature not enabled on old SMART3G & NC) | 780.84

## SMART S-MS MED
Auto+Hide+Values
SMART S-MS MED Marine Approved Gas Detector
The SMART S-MS MED line is designed to withstand the harshest marine environments, enabling the monitoring of flammable gases using either catalytic or infrared technologies, and refrigerant gases with an infrared sensor. These detectors are certified to SIL 2 for hardware and SIL 3 for software capability, and are ATEX certified for Zone 1 as II2G Ex db IIC T5 Gb. Additionally, the SMART S-MS line has received marine certification from Lloyd's Register.
Product Configurator
| Transmitter | SMS3A11MED | Electronic with1 output RS485+4-20 mA | 869.94
| Sensor Head | S4032R1234YF | Smart S Sensor Mod, R1234yf IR 0-2000 ppm | 2587.14
Use the configurator to generate part number, price and description.
| Configured Code No. |  | PVP
| SMS3A11MED-S4032R1234YF |  | 5601.96
Electronic with1 output RS485+4-20 mA - Smart S Sensor Mod, R1234yf IR 0-2000 ppm
Smart S-MS MED Transmitters 3A Aluminum Enclosure complete with electronics
| Product Code | Description | PVP
| SMS3A11MED | Electronic with1 output RS485+4-20 mA | 869.94
| SMS3A13MED | Electronic with1 output RS485+4-20 mA, relay card | 1028.7
Smart S-MS MED Sensor Modules
Catalitic Sensor - Flammable Gases
| Product Code | Description | PVP | Gas Cylinder
| S4003ME | Smart S Sensor Mod, Methane Cat 0-100% LEL | 928.26 | SIB09
| S4003PR | Smart S Sensor Mod, Propane Cat 0-100% LEL | 928.26 | SIB06
| S4003BU | Smart S Sensor Mod, Butane Cat 0-100% LEL | 928.26 | SIB06
| S4003PE | Smart S Sensor Mod, Pentane Cat 0-100% LEL | 928.26 | SIB06
| S4003H2 | Smart S Sensor Mod, Pentane Cat 0-100% LEL | 928.26 | SIB19
| S4003VB | Smart S Sensor Mod, Petrol Vapours Cat 0-100% LEL | 928.26 | SIB06
Infrared Sensor - Flammable Gases
| Product Code | Description | PVP | Gas Cylinder
| S4019ME | Smart S Sensor Mod, Methane Cat 0-100% LEL | 1438.56 | SIB09
| S4004PR | Smart S Sensor Mod, Propane IR 0-100% LEL | 1438.56 | SIB28
| S4004BU | Smart S Sensor Mod, Butane IR 0-100% LEL | 1438.56 | SIB28
| S4004PE | Smart S Sensor Mod, Pentane IR 0-100% LEL | 1438.56 | SIB28
Infrared Sensor - A1, A2L Refrigerant Gases
| Product Code | Description | PVP | Gas Cylinder
| S4152R32 | Smart S Sensor Mod, R32 IR 0-2000 ppm | 2587.14 | RIB148-34
| S4026R134A | Smart S Sensor Mod, R134a IR 0-2000 ppm | 2587.14 | RIB144-34
| S4031R125 | Smart S Sensor Mod, R125 IR 0-2000 ppm | 2587.14 | RIB130-34
| S4032R1234YF | Smart S Sensor Mod, R1234yf IR 0-2000 ppm | 2587.14 | RIB145-34
| S4153R1234ZE | Smart S Sensor Mod, R1234ze IR 0-2000 ppm | 2587.14 | RIB143-34
| S4033R404A | Smart S Sensor Mod, R404a IR 0-2000 ppm | 2587.14 | RIB131-34
| S4034R407A | Smart S Sensor Mod, R407a IR 0-2000 ppm | 2587.14 | RIB132-34
| S4034R407F | Smart S Sensor Mod, R407f IR 0-2000 ppm | 2587.14 | RIB133-34
| S4154R449A | Smart S Sensor Mod, R449a IR 0-2000 ppm | 2587.14 | RIB134-34
| S4035R507A | Smart S Sensor Mod, R507a IR 0-2000 ppm | 2587.14 | RIB135-34
| S4036R410A | Smart S Sensor Mod, R410a IR 0-2000 ppm | 2587.14 | RIB147-34
| S4155R417A | Smart S Sensor Mod, R417a IR 0-2000 ppm | 2587.14 | RIB136-34

## SMART S-IS
Auto+Hide+Values
SMART S-IS Intrinsically Safe Gas Detector
The SMART S-IS line is engineered to meet the most demanding industrial requirements, providing reliable monitoring of toxic gases in harsh environments. These detectors feature a Zener barrier and galvanic separation, and are ATEX certified as II2G Ex db [ib] ib IIC T6 Gb.
Product Configurator
| Transmitter | SIS4A21 | Sens Smart S-IS Transmitter, 4A Alum Encl, RS-485, 4-20 mA, LCD Disp, IS&GI card | 1739.88 | Sens Smart S-IS Transmitter, 4A Alum Encl, RS-485, 4-20 mA, LCD Disp, IS&GI card
| Sensor | S5005O2 | Smart S Sensor Mod, Oxygen O2 EC 0-25% Vol. | 1135.62 | Smart S Sensor Mod, Oxygen O2 EC 0-25% Vol.
Use the configurator to generate part number, price and description:
| Configured Part No. |  | PVP
| SIS4A21-S5005O2 |  | 2875.5
Sens Smart S-IS Transmitter, 4A Alum Encl, RS-485, 4-20 mA, LCD Disp, IS&GI card Smart S Sensor Mod, Oxygen O2 EC 0-25% Vol.
Smart S-IS Transmitters
Smart S-IS Transmitters 4A Aluminum Enclosure complete with electronics
| Product Code | Description | PVP
| SIS4A21 | Sens Smart S-IS Transmitter, 4A Alum Encl, RS-485, 4-20 mA, LCD Disp, IS&GI card | 1739.88
| SIS4A22 | Sens Smart S-IS Transmitter, 4A Alum Encl, RS-485, 4-20 mA, LED Disp, IS&GI card | 2055.78
| SIS4A23 | Sens Smart S-IS Transmitter , 4A Alum Encl, RS-485, 4-20 mA, LCD Disp, IS&GI card, 3x Relay | 1848.42
| SIS4A24 | Sens Smart S-IS Transmitter , 4A Alum Encl, RS-485, 4-20 mA, LED Disp, IS&GI card, 3x Relay | 2151.36
Smart S-IS Transmitters 2S Stainless Steel Enclosure complete with electronics
| Product Code | Description | PVP
| SIS2S21 | Sens Smart S-IS Transmitter, 2S SS Encl, RS-485, 4-20 mA, LCD Disp, IS&GI card | 2366.82
| SIS2S22 | Sens Smart S-IS Transmitter, 2S SS Encl, RS-485, 4-20 mA, LED Disp, IS&GI card | 2669.76
| SIS2S23 | Sens Smart S-IS Transmitter, 2S SS Encl, RS-485, 4-20 mA, LCD display, IS&GI card, 3x Relay | 2465.64
| SIS2S24 | Sens Smart S-IS Transmitter, 2S SS Encl, RS-485, 4-20 mA, LED display, IS&GI card, 3x Relay | 2776.68
Smart S-IS Sensor Modules
| Product Code | Description | PVP | Gas Cylinder
| S5005O2 | Smart S Sensor Mod, Oxygen O2 EC 0-25% Vol. | 1135.62 | SIB62
| S5011SD | Smart S Sensor Mod, Sulphur Dioxide SO2 EC 0-20 ppm | 1547.1 | RIB91-34
| S5049SD | Smart S Sensor Mod, Sulphur Dioxide S02 EC 0-100 ppm | 1547.1 | RIB129-34
| S5012ND | Smart S Sensor Mod, Nitrogen Dioxide NO2 EC 0-20 ppm | 1547.1 | RIB116-34
| S5015CL | Smart S Sensor Mod, Chlorine Cl2 EC 0-20 ppm | 1547.1 | RIB116-34
| S5014HC | Smart S Sensor Mod, Hydrochloric Acid HCl EC 0-50 ppm | 1547.1 | RIB82-34

## Dati (hidden)
SSS3A11 | SMS3A11 | SMS3A11MED | SIS4A21
SSS3A12 | SMS3A13 | SMS3A13MED | SIS4A22
SSS3A13 | SMS3A17 |  | SIS4A23
SSS3A14 | SMS4A11 |  | SIS4A24
SSS3A17 | SMS4A13 |  | SIS2S21
SSS3A18 | SMS4A17 |  | SIS2S22
SSS4A11 | SMS2S11 |  | SIS2S23
SSS4A12 | SMS2S13 |  | SIS2S24
SSS4A13 | SMS2S17
SSS4A14
SSS4A17
SSS4A18
SSS2S11
SSS2S12
SSS2S13
SSS2S14
SSS2S17
SSS2S18
S4003ME | S4003ME | S4152R32 | S5005O2
S4001ME | S4001ME | S4026R134A | S5011SD
S4003PR | S4003PR | S4031R125 | S5049SD
S4003BU | S4003BU | S4032R1234YF | S5012ND
S4003VB | S4003VB | S4032R1234ZE | S5015CL
S4003AT | S4003AT | S4033R404A | S5014HC
S4037AC | S4037AC | S4034R407A
S4003AM | S4003AM | S4034R407F
S4003CP | S4003CP | S4154R449A
S4003EN | S4003EN | S4035R507A
S4003AE | S4003AE | S4036R410A
S4003ET | S4003ET | S4155R417A
S4003EL | S4003EL
S4003EP | S4003EP
S4003ES | S4003ES
S4003H2 | S4003H2
S4003IB | S4003IB
S4002P8 | S4002P8
S4003MT | S4003MT
S4003MK | S4003MK
S4003PE | S4003PE
S4003PP | S4003PP
S4003PN | S4003PN
S4003ST | S4003ST
S4003TO | S4003TO
S4003XI | S4003XI
S4019ME | S4019ME
S4004PR | S4001ME
S4004BU | S4004PR
S4004AT | S4004BU
S4004BN | S4004AT
S4004CE | S4004BN
S4004DT | S4004CE
S4004EN | S4004DT
S4004AE | S4004EN
S4004ET | S4004AE
S4004EP | S4004ET
S4004ES | S4004EP
S4004IB | S4004ES
S4004IL | S4004IB
S4004IP | S4004IL
S4004PE | S4004IP
S4004PP | S4004PE
S4004TO | S4004PP
S4004DE | S4004TO
S4157R454B | S4004DE
S4065R32 | S4157R454B
S4066R1234YF | S4065R32
S4067R1234ZE | S4066R1234YF
S4152R32 | S4067R1234ZE
S4026R134A | S4152R32
S4027SF6 | S4026R134A
S4031R125 | S4027SF6
S4032R1234YF | S4031R125
S4153R1234ZE | S4032R1234YF
S4033R404A | S4153R1234ZE
S4034R407A | S4033R404A
S4034R407F | S4034R407A
S4154R449A | S4034R407F
S4035R507A | S4154R449A
S4036R410A | S4035R507A
S4155R417A | S4036R410A
S4156R448 | S4155R417A
S4015CO2 | S4156R448
S4016CO2 | S4015CO2
S4017CO2 | S4016CO2
S4005O2 | S4017CO2
S4166CO | S4005O2
S4006CO | S4166CO
S4021CO | S4006CO
S4022AM | S4021CO
S4007AM | S4022AM
S4008AM | S4007AM
S4009HS | S4008AM
S4028HS | S4009HS
S4029HS | S4028HS
S4010NO | S4029HS
S4089H2 | S4010NO
S4168H2 | S4089H2
S4169H2 | S4168H2
S4167IL | S4169H2
S4158IL | S4167IL
S4159IL | S4158IL
S4075BZ | S4159IL
S4160BZ | S4075BZ
S4161BZ | S4160BZ
S4075TO | S4161BZ
S4160TO | S4075TO
S4161TO | S4160TO
S4161TO

## COVER SMART 3
#VALUE!

## SMART 3-R
Auto+Hide+Values
SMART 3-R Gas detector for machinery rooms, laboratories & safe areas
The SMART 3-R detectors are designed for non-classified areas, making them ideal for detecting flammable, toxic, and refrigerant gases in non-hazardous environments. They offer analog 4-20 mA and digital RS485 transmissions, along with 3-relay outputs. A multi-color LED light ring provides a visual indication of the detector's status, while an internal buzzer alerts users to any fault or alarm condition. Rated IP42, the SMART 3-R features one-man calibration using the STS/CKD series calibration keypad. Additionally, they can be supplied with an optional remote sensor head and come with a 5-meter cable. Compliance: • EN 378-3 • EN14624 • EN 50676 • IEC 60335-2-40 Annex LL • ISO 5149-3 • F-Gas Regulation Features: • Standard 4-20 mA, RS485 and 3-relay outputs • Multicolour LED ring • Internal buzzer for audible warning • Ongoing system self-diagnosis • Optional remote sensor head • Standard 12-24 Vdc power supply. Optional 90-240 Vac, without 4-20 mA output • IP42
Flammable Catalytic Sensors
| Product Code | Description | PVP | Sensor Head | Gas Cylinder
| S4590ME | Smart 3-R, Methane CH4 Cat. 0-100% LEL, 4-20 mA | 819.72 | K4590ME | SIB09
| S4590GP | Smart 3-R, LPG Cat. 0-100% LEL, 4-20 mA | 819.72 | K4590GP | SIB06
| S4590PR | Smart 3-R, Propane Cat. 0-100% LEL, 4-20 mA | 819.72 | K4590PR | SIB09
| S4590BU | Smart 3-R, Butane Cat. 0-100% LEL, 4-20 mA | 819.72 | K4590BU | SIB06
| S4590VB | Smart 3-R, Petrol Vapour Cat. 0-100% LEL, 4-20 mA | 819.72 | K4590VB | SIB06
Flammable Infra-Red (IR) Sensors
| Product Code | Description | PVP | Sensor Head | Gas Cylinder
| S4591ME | Smart 3-R, Methane CH4 IR 0-100% LEL, 4-20 mA | 1466.1 | K4591ME | SIB09
| S4592PR | Smart 3-R, Propane IR 0-100% LEL, 4-20 mA | 1466.1 | K4592PR | SIB28
| S4592BU | Smart 3-R, Butane IR 0-100% LEL, 4-20 mA | 1466.1 | K4592BU | SIB28
MPS Sensors
| Product Code | Description | PVP | Sensor Head | Gas Cylinder
| S4949ME | Smart 3-R, Methane CH4 MPS 0-100% LEL, 4-20 mA | 2174.04 | K4949ME | SIB09
| S4949H2 | Smart 3-R, Hydrogen H2 MPS 0-100% LEL, 4-20 mA | 2174.04 | K4949H2 | SIB19
Toxic Electrochemical Sensors and Oxygen
| Product Code | Description | PVP | Sensor Head | Gas Cylinder
| S4593O2 | Smart 3-R, Oxygen O2 EC 0-25% Vol., 4-20 mA | 1048.14 | K4593O2 | SIB62
| S4595CO | Smart 3-R, Carbon Monoxide CO EC 0-500 ppm, 4-20 mA | 1048.14 | K4595CO | SIB56
| S4597AM | Smart 3-R, Ammonia NH3 EC 0-100 ppm, 4-20 mA | 1466.1 | K4597AM | RIB73-34
| S4598AM | Smart 3-R, Ammonia NH3 EC 0-200 ppm, 4-20 mA | 1466.1 | K4598AM | RIB74-34
| S4599AM | Smart 3-R, Ammonia NH3 EC 0-1000 ppm, 4-20 mA | 1466.1 | K4599AM | RIB76-34
| S4634CL | Smart 3-R, Chlorine Cl2 EC 0-20 ppm, 4-20 mA | 1722.06 | K4634CL | RIB116-34
| S4601HS | Smart 3-R, Hydrogen Sulphide H2S EC 0-50 ppm, 4-20 mA | 1722.06 | K4601HS | RIB81-34
| S4630SD | Smart 3-R, Sulphur Dioxide SO2 EC 0-20 ppm, 4-20 mA | 1722.06 | K4630SD | RIB91-34
| S4632ND | Smart 3-R, Nitrogen Dioxide NO2 EC 0-20 ppm, 4-20 mA | 1722.06 | K4632ND | RIB116-34
CO2 Infra-Red (IR) Sensors
| Product Code | Description | PVP | Sensor Head | Gas Cylinder
| S4606CO2 | Smart 3-R, Carbon Dioxide CO2 IR 0-2% Vol., 4-20 mA | 1537.38 | K4606CO2 | SIB42
| S4607CO2 | Smart 3-R, Carbon Dioxide CO2 IR 0-5% Vol., 4-20 mA | 1537.38 | K4607CO2 | SIB44
| S4608CO2 | Smart 3-R, Carbon Dioxide CO2 IR 0-30% Vol., 4-20 mA | 1537.38 | K4608CO2 | SIB48
| S4610CO2 | Smart 3-R, Carbon Dioxide CO2 IR 0-5000ppm, 4-20 mA | 1537.38 | K4610CO2 | RIB110-34
Refrigerant Infra-Red (IR) Sensors
| Product Code | Description | PVP | Sensor Head | Gas Cylinder
| S4592DE | Smart 3-R, R152a IR 0-100% LEL, 4-20 mA | 1466.1 | K4592DE | SIB26
| S4629R454B | Smart 3-R, R454b IR 0-100% LEL, 4-20 mA | 2114.1 | K4629R454B | SIB507
| S4626R32 | Smart 3-R, R32 IR 0-100% LEL detector, 4-20 mA | 2114.1 | K4626R32 | SIB512
| S4627R1234YF | Smart 3-R, R1234yf IR 0-100% LEL, 4-20 mA | 2114.1 | K4627R1234YF | SIB501
| S4628R1234ZE | Smart 3-R, R1234ze IR 0-100% LEL, 4-20 mA | 2114.1 | K4628R1234ZE | SIB502
| S4638R32 | Smart 3-R, R32 IR 0-5000ppm, 4-20 mA | 1479.06 | K4638R32 | RIB148-34
| S4639R134A | Smart 3-R, R134a IR 0-5000ppm, 4-20 mA | 1479.06 | K4639R134A | RIB144-34
| S4640R404A | Smart 3-R, R404a IR 0-5000ppm, 4-20 mA | 1479.06 | K4640R404A | RIB131-34
| S4641R407C | Smart 3-R, R407c IR 0-5000ppm, 4-20 mA | 1479.06 | K4641R407C | RIB151-34
| S4642R410A | Smart 3-R, R410a IR 0-5000ppm, 4-20 mA | 1479.06 | K4642R410A | RIB147-34
| S4643R1234YF | Smart 3-R, R1234yf IR 0-5000ppm, 4-20 mA | 1479.06 | K4643R1234YF | RIB145-34
| S4644R1234ZE | Smart 3-R, R1234ze IR 0-5000ppm, 4-20 mA | 1479.06 | K4644R1234ZE | RIB143-34
| S4857R1233ZD | Smart 3-R, R1233Zd IR 0-5000ppm, 4-20 mA | 1479.06 | K4857R1233ZD | RIB138-34
| S4611R32 | Smart 3-R, R32 IR 0-2000ppm, 4-20 mA | 2114.1 | K4611R32 | RIB148-34
| S4612R134A | Smart 3-R, R134a IR 0-2000ppm, 4-20 mA | 2114.1 | K4612R134A | RIB144-34
| S4613SF6 | Smart 3-R, SF6 IR 0-2000ppm, 4-20 mA | 2114.1 | K4613SF6 | RIB149-34
| S4614R125 | Smart 3-R, R125 IR 0-2000ppm, 4-20 mA | 2114.1 | K4614R125 | RIB130-34
| S4615R1234YF | Smart 3-R, R1234yf IR 0-2000ppm, 4-20 mA | 2114.1 | K4615R1234YF | RIB145-34
| S4616R1234ZE | Smart 3-R, R1234ze IR 0-2000ppm, 4-20 mA | 2114.1 | K4616R1234ZE | RIB143-34
| S4617R404A | Smart 3-R, R404a IR 0-2000ppm, 4-20 mA | 2114.1 | K4617R404A | RIB131-34
| S4618R407A | Smart 3-R, R407a IR 0-2000ppm, 4-20 mA | 2114.1 | K4618R407A | RIB132-34
| S4619R407F | Smart 3-R, R407f IR 0-2000ppm, 4-20 mA | 2114.1 | K4619R407F | RIB133-34
| S4620R449A | Smart 3-R, R449a IR 0-2000ppm, 4-20 mA | 2114.1 | K4620R449A | RIB134-34
| S4621R507A | Smart 3-R, R507a IR 0-2000ppm, 4-20 mA | 2114.1 | K4621R507A | RIB135-34
| S4622R410A | Smart 3-R, R410a IR 0-2000ppm, 4-20 mA | 2114.1 | K4622R410A | RIB147-34
| S4623R417A | Smart 3-R, R417a IR 0-2000ppm, 4-20 mA | 2114.1 | K4623R417A | RIB136-34
| S4624R448 | Smart 3-R, R448 IR 0-2000ppm, 4-20 mA | 2114.1 | K4624R448 | RIB137-34
SMART3-R can be supplied with: - 90-240 Vac power supply (no 4-20 mA transmission) - Remote sensor head Add extension “-220” to the P/N to get the detector at 90-240 Vac, as follows, at no extra cost. Add extension “-RH” to the P/N to get the detector with remote sensor head at 12-24 Vdc Add extension “-220-RH” to the P/N to get the detector with remote sensor head and 90-240 Vac power supply. Example: S4590ME-220 S4590ME-RH S4590ME-220-RH
Additional costs to be added to the list price for remote sensor heads
| Product Code | Description | List Price
| -RH | With Remote Sensor Head | 125
| -220-RH | With Remote Sensor Head and 90-240 Vac Power Supply | 125
Accessories
| Product Code | Description | List Price
| STS/CKD+ | Enhanced calibration keypad with 7-digit display and SD card slot. Compatible with Smart series for operations like zero/span calibration and monitor 4-20 mA signals. | 318
| STS/CKD-OLED | Enhanced calibration keypad with OLED display. It can be connected to all SENSITRON’S SMART detectors to adjust the Zero, Span and 4-20 mA values. This version offers the visualization of the instantaneous concentration in graphic format. | 482

## SMART 3 BLIZZARD
Auto+Hide+Values
SMART 3 BLIZZARD Gas detector for machine rooms, laboratories, and safe areas
The SMART 3 BLIZZARD detector enables the detection of refrigerant gases in safe (non-classified) areas. The detector provides a 4-20 mA output, 2 relays, and a wireless interface. The circular LED indicator on the front panel displays the instrument’s operating status, while an internal buzzer is activated in the event of a fault and/or alarm. The SMART 3 BLIZZARD detector is ideal for applications in machine rooms, laboratories, and non-classified areas. Instrument control and calibration are carried out via a wireless connection and the dedicated SENSE interface, which can be accessed from a PC or mobile device using standard web browsers. The detector has an IP54 protection rating. SMART 3 BLIZZARD complies with the following standards: Compliance: • EN 378-3 • EN14624 - in progress • EN 50676 - in progress • IEC 60335-2-40 Annex LL (sostituita da TS 63542) - in progress • ANSI/ASHRAE 15 • CSA B52: 2023 • F-Gas Regulation EU 573/2024 • Certificato SIL2 secondo EN50402 – in progress Features: •Non-intrusive field calibration via SENSE interface •Infrared sensor •4-20 mA output, 2 relays, and wireless interface •Internal buzzer and high-brightness smart LEDs •Continuous self-diagnostics •Easy sensor replacement •12-24 VDC power supply
Refrigerant Infra-Red (IR) Sensors A1 & A2L
| Product Code | Description | PVP | Sensor Head | Gas Cylinder
| S5116R32 | SMART 3 Blizzard per R32, 0 - 100% LFL, 4-20mA | 1140.48 | K5116R32 | SIB512
| S5117R454B | SMART 3 Blizzard per R454B, 0 - 100% LFL, 4-20mA | 1140.48 | K5117R454B | SIB507
| S5118R410A | SMART 3 Blizzard per R410A, 0 - 10.000 ppm, 4-20mA | 1140.48 | K5118R410A | RIB147-34
| S5226R1234YF | SMART 3 Blizzard per R1234YF, 0 - 5.000 ppm, 4-20mA | 1354.32 | K5226R1234YF | RIB145-34
| S5227R1234ZE | SMART 3 Blizzard per R1234ZE, 0 - 5.000 ppm, 4-20mA | 1354.32 | K5227R1234ZE | RIB143-34
| S5228R134A | SMART 3 Blizzard per R134A, 0 - 5.000 ppm, 4-20mA | 1354.32 | K5228R134A | RIB144-34
| S5229R404A | SMART 3 Blizzard per R404A, 0 - 5.000 ppm, 4-20mA | 1354.32 | K5229R404A | RIB131-34

## SMART 3 NC
Auto+Hide+Values | SMART 3 NC Gas detector for safe areas
The SMART 3-NC detectors are designed to offer cost-effective solutions with high-tech professional features for car parks and non-classified areas. They are available for detecting Methane, LPG, Petrol Vapours, CO, NO2, and CO2. These detectors use a professional catalytic sensor for flammable gas detection, an electrochemical cell for CO, and an infrared sensor for CO2 detection. The SMART 3-NC provides an analog 4-20 mA output signal and can be upgraded with additional features to include RS485 communication or a three-relay output. These gas detectors are rated IP55 for protection against dust and water. Features: • Standard 4-20 mA 3 wire output • Optional 3 relay output • Optional RS-485 interface for Modbus communication. • Ongoing system self diagnosis • IP55
For Car Park Applications
| Product Code | Description | PVP | Sensor Head | Gas Cylinder
| S1450CO | Smart 3-NC, Carbon Monoxide CO EC 0-300 ppm, 4-20 mA | 520.02 | K1450CO | SIB56
| S1455VB | Smart 3-NC, Petrol Vapour Cat. 0-100% LEL, 4-20 mA | 557.28 | K1455VB | SIB06
| S1451ND | Smart 3-NC, Nitrogen Dioxide NO2 EC 0-300 ppm, 4-20 mA | 756.54 | K1451ND | RIB116-34
Flammable Gas Detectors for Safe Area
| Product Code | Description | PVP | Sensor Head | Gas Cylinder
| S1455ME | Smart 3-NC, Methane CH4 Cat. 0-100% LEL, 4-20 mA | 557.28 | K1455ME | SIB09
| S1455GP | Smart 3-NC, LPG Cat. 0-100% LEL, 4-20 mA | 557.28 | K1455GP | SIB06
| S1455PR | Smart 3-NC, Propane Cat. 0-100% LEL, 4-20 mA | 557.28 | K1455PR | SIB09
Carbon Dioxide Gas Detector
| Product Code | Description | PVP | Sensor Head | Gas Cylinder
| S2401CO2 | Smart 3-NC, Carbon Dioxide CO2 IR 0-5000 ppm, 4-20 mA | 1394.82 | K2401CO2 | RIB110-34
Accessories
| Product Code | Description | PVP
| STS/CSW+ | Calibration software for SMART3G, SMART3NC & SMART3R detectors consisting of a USB key, a USB/RS485 converter and the STS/IDI card. The software runs on Windows XP, Windows 7 e Windows 10. The calibration software allows to modify the alarm thresholds and the measuring range, and allows adjusting the Zero, Span and 4-20 mA values via a user-friendly menu, either in situ or in laboratory. | 626.94
| STS/IDI | Plug-in RS485 interface, to make SMART3NC series gas detectors communicate in Modbus protocol with MULTISCAN++S1, MULTISCAN 8+ & communicate in Modbus protocol with MULTISCAN++S1, MULTISCAN 8+ & STPL4+D control panels. | 84.24
| STS/CKD+ | Enhanced calibration keypad with 7-digit display and SD card slot. Compatible with Smart series for operations like zero/span calibration and monitor 4-20 mA signals. | 515.16
| STS/CKD-OLED | Enhanced calibration keypad with OLED display and SD-CARD slot. It can be connected to all SMART detectors to adjust the Zero, Span and 4-20 mA values. This version offers enhanced features like: - visualization of the instantaneous concentration in graphic format; - possibility to monitor the detector trend and download the data stored in the SD-CARD - detector’s firmware update via SD-CARD (feature not enabled on old SMART3G and NC detector) | 780.84
| STS3REL | 3-Relay card for SMART3 NC gas detectors. Relay at 12÷24V | 124.74

## SMART 3H Lite
Auto+Hide+Values
SMART 3H LITE A1 and A2L refrigerant Gas Detector
The SMART 3H LITE is designed to meet cost-effective requirements, enabling the monitoring of A1 and A2L refrigerant gases in hotel and building applications. It features three mode status LEDs for visual indication of the detector's status, and an internal buzzer that activates to warn of any fault or alarm condition. The detectors use a unique, cost-effective IR sensor and provide RS485 communication along with two pre-set relay outputs. Compliance: • EN 378-3 • EN14624 • EN 50676 • IEC 60335-2-40 Annex LL • ISO 5149-3 • F-Gas Regulation Features: • 2 pre-set output relays • RS485 communication (0-20%LFL and 0-5000 ppm) • Multicolour LED ring • Standard ranges: 0-20% LFL and 0-5000 ppm • Internal buzzer for audible warning • Ongoing system self-diagnosis • Standard 12-24 Vdc power supply
Refrigerant Infra-Red (IR) Sensors with 0-20% LEL range
| Product Code | Description | PVP | Gas Cylinder
| S3843R32 | Smart 3H Lite, R32 IR 0-20% LEL, AL1 5% LEL, AL2 10% LEL | 1153.44 | SIB503
| S3844R1234YF | Smart 3H Lite, R1234yf IR 0-20% LEL, AL1 5% LEL, AL2 10% LEL | 1153.44 | SIB504
| S3845R1234ZE | Smart 3H Lite, R1234ze IR 0-20% LEL, AL1 5% LEL, AL2 10% LEL | 1153.44 | SIB505
| S3848R454A | Smart 3H Lite, R454A IR 0-20% LEL, AL1 5% LEL, AL2 10% LEL | 1153.44 | On request
| S3849R454B | Smart 3H Lite, R454B IR 0-20% LEL, AL1 5% LEL, AL2 10% LEL | 1153.44 | SIB507
| S3850R454C | Smart 3H Lite, R454C IR 0-20% LEL, AL1 5% LEL, AL2 10% LEL | 1153.44 | On request
| S3851R452B | Smart 3H Lite, R452B IR 0-20% LEL, AL1 5% LEL, AL2 10% LEL | 1153.44 | On request
Refrigerant Infra-Red (IR) Sensors with 0-5000 ppm range
| Product Code | Description | PVP | Gas Cylinder
| S3846R134A | Smart 3H Lite, R134a IR 0-5000 ppm, AL1 1000 ppm, AL2 2000 ppm | 1140.48 | RIB144-34
| S3847R410A | Smart 3H Lite, R410A IR 0-5000 ppm, AL1 1000 ppm, AL2 2000 ppm | 1140.48 | RIB147-34
| S3854R32 | Smart 3H Lite, R32 IR 0-5000 ppm, AL1 1000 ppm, AL2 2000 ppm | 1140.48 | RIB148-34
| S3855R404A | Smart 3H Lite, R404A IR 0-5000 ppm, AL1 1000 ppm, AL2 2000 ppm | 1140.48 | RIB131-34
| S3856R407C | Smart 3H Lite, R407C IR 0-5000 ppm, AL1 1000 ppm, AL2 2000 ppm | 1140.48 | RIB151-34
| S3857R1234YF | Smart 3H Lite, R1234YF IR 0-5000 ppm, AL1 1000 ppm, AL2 2000 ppm | 1140.48 | RIB145-34
| S3858R1234ZE | Smart 3H Lite, R1234E IR 0-5000 ppm, AL1 1000 ppm, AL2 2000 ppm | 1140.48 | RIB143-34
| S3859CO2 | Smart 3H Lite, CO2 IR 0-5000 ppm, AL1 1000 ppm, AL2 2000 ppm | 1140.48 | SIB40
Accessories
| Product Code | Description | PVP
| STS/CKD+ | Enhanced calibration keypad with 7-digit display and SD card slot. Compatible with Smart series for operations like zero/span calibration and monitor 4-20 mA signals. | 515.16

## SMART 3H FM-LITE
Auto+Hide+Values
SMART 3H FM-LITE
The SMART 3H FM-LITE have been designed to offer a flush mount solution for refrigerant gases to be detected in hotels or residential buildings. The detectors employ a unique cost-effective IR sensor and provide RS485 communication and two pre-set relay outputs. Compliance: • EN 378-3 • IEC 60335-2-40 Annex LL • EN14624 • EN 50676 • ISO 5149-3 • F-Gas Regulation Features: • standard flush-mounting box 3 modules • 2 output relays • RS485 communication (0-20%LFL and 0-5000 ppm) • Standard ranges: 0-20% LFL and 0-5000 ppm • LED indication & internal buzzer for audible warning • Ongoing system self-diagnosis • Standard 90-240 Vac or 12-24 Vdc • IP42
90-240V AC Power Input
| Product Code | Description | PVP | Gas Cylinder
| S4143R32 | Smart 3H FM-LITE, VAC, R32 IR 0-20% LEL, AL1 5% LEL, AL2 10% LEL | 1038.42 | SIB503
| S4144R1234YF | Smart 3H FM-LITE, VAC, R1234yf IR 0-20% LEL, AL1 5% LEL, AL2 10% LEL | 1038.42 | SIB504
| S4145R1234ZE | Smart 3H FM-LITE, VAC, R1234ze IR 0-20% LEL, AL1 5% LEL, AL2 10% LEL | 1038.42 | SIB505
| S4660R454A | Smart 3H FM-LITE, VAC, R454a IR 0-20% LEL, AL1 5% LEL, AL2 10% LEL | 1038.42 | On request
| S4661R454B | Smart 3H FM-LITE, VAC, R454b IR 0-20% LEL, AL1 5% LEL, AL2 10% LEL | 1038.42 | SIB507
| S4662R454C | Smart 3H FM-LITE, VAC, R454c IR 0-20% LEL, AL1 5% LEL, AL2 10% LEL | 1038.42 | On request
| S4663R452B | Smart 3H FM-LITE, VAC, R452b IR 0-20% LEL, AL1 5% LEL, AL2 10% LEL | 1038.42 | On request
90-240V AC Power Input
| Product Code | Description | PVP | Gas Cylinder
| S4146R134A | Smart 3H FM-LITE, VAC, R134a IR 0-5000 ppm, AL1 1000 ppm, AL2 2000 ppm | 1038.42 | RIB144-34
| S4147R410A | Smart 3H FM-LITE, VAC, R410a IR 0-5000 ppm, AL1 1000 ppm, AL2 2000 ppm | 1038.42 | RIB147-34
| S4669R32 | Smart 3H FM-LITE, VAC, R32 IR 0-5000 ppm, AL1 1000 ppm, AL2 2000 ppm | 1038.42 | RIB148-34
| S4670R404A | Smart 3H FM-LITE, VAC, R404a IR 0-5000 ppm, AL1 1000 ppm, AL2 2000 ppm | 1038.42 | RIB131-34
| S4671R407C | Smart 3H FM-LITE, VAC, R407c IR 0-5000 ppm, AL1 1000 ppm, AL2 2000 ppm | 1038.42 | RIB151-34
| S4672R1234YF | Smart 3H FM-LITE, VAC, R1234yf IR 0-5000 ppm, AL1 1000 ppm, AL2 2000 ppm | 1038.42 | RIB145-34
| S4673R1234ZE | Smart 3H FM-LITE, VAC, R1234e IR 0-5000 ppm, AL1 1000 ppm, AL2 2000 ppm | 1038.42 | RIB143-34
| S4674CO2 | Smart 3H FM-LITE, VAC, Carbon Dioxide CO2 IR 0-5000 ppm, AL1 1000 ppm, AL2 2000 ppm | 1038.42 | SIB40
12-24V DC Power Input
| Product Code | Description | PVP | Gas Cylinder
| S4138R32 | Smart 3H FM-LITE, VDC, R32 IR 0-20% LEL, AL1 5% LEL, AL2 10% LEL | 1038.42 | SIB503
| S4139R1234YF | Smart 3H FM-LITE, VDC, R1234yf IR 0-20% LEL, AL1 5% LEL, AL2 10% LEL | 1038.42 | SIB504
| S4140R1234ZE | Smart 3H FM-LITE, VDC, R1234ze IR 0-20% LEL, AL1 5% LEL, AL2 10% LEL | 1038.42 | SIB505
| S4680R454A | Smart 3H FM-LITE, VDC, R454a IR 0-20% LEL, AL1 5% LEL, AL2 10% LEL | 1038.42 | On request
| S4681R454B | Smart 3H FM-LITE, VDC, R454b IR 0-20% LEL, AL1 5% LEL, AL2 10% LEL | 1038.42 | SIB507
| S4682R454C | Smart 3H FM-LITE, VDC, R454c IR 0-20% LEL, AL1 5% LEL, AL2 10% LEL | 1038.42 | On request
| S4683R452B | Smart 3H FM-LITE, VDC, R452b IR 0-20% LEL, AL1 5% LEL, AL2 10% LEL | 1038.42 | On request
12-24V DC Power Input
| Product Code | Description | PVP | Gas Cylinder
| S4141R134A | Smart 3H FM-LITE, VDC, R134a IR 0-5000 ppm, AL1 1000 ppm, AL2 2000 ppm | 1038.42 | RIB144-34
| S4142R410A | Smart 3H FM-LITE, VDC, R410a IR 0-5000 ppm, AL1 1000 ppm, AL2 2000 ppm | 1038.42 | RIB147-34
| S4689R32 | Smart 3H FM-LITE, VDC, R32 IR 0-5000 ppm, AL1 1000 ppm, AL2 2000 ppm | 1038.42 | RIB148-34
| S4690R404A | Smart 3H FM-LITE, VDC, R404a IR 0-5000 ppm, AL1 1000 ppm, AL2 2000 ppm | 1038.42 | RIB131-34
| S4691R407C | Smart 3H FM-LITE, VDC, R407c IR 0-5000 ppm, AL1 1000 ppm, AL2 2000 ppm | 1038.42 | RIB151-34
| S4692R1234YF | Smart 3H FM-LITE, VDC, R1234yf IR 0-5000 ppm, AL1 1000 ppm, AL2 2000 ppm | 1038.42 | RIB145-34
| S4693R1234ZE | Smart 3H FM-LITE, VDC, R1234ze IR 0-5000 ppm, AL1 1000 ppm, AL2 2000 ppm | 1038.42 | RIB143-34
| S4694CO2 | Smart 3H FM-LITE, VDC, Carbon Dioxide CO2 IR 0-5000 ppm, AL1 1000 ppm, AL2 2000 ppm | 1038.42 | SIB40
Accessories
| Product Code | Description | PVP
| STS/CKD+ | Enhanced calibration keypad with 7-digit display. It can be connected to Smart detectors to adjust the Zero, Span and 4-20 mA values | 515.16

## COVER CONTROL PANELS
#VALUE!

## PL4+ & PL4+D
Auto+Hide+Values
PL4+ and PL4+D
The PL4+ and PL4+D are advanced gas control units designed to ensure safety and reliability in monitoring flammable, toxic, and oxygen gas levels in various environments. Both units offer precise gas concentration measurements and alarm set points through their LCD displays, with robust features tailored to different installation needs. Certified to SIL1 and ATEX standards (EN60079-29-1), these units provide dependable performance and compliance with stringent safety regulations. - SIL1 and ATEX certified according to EN60079-29-1, ensuring high safety and compliance standards. - Features an LCD display that provides precise indications of gas concentrations and alarm set points. - Monitor multiple gas detectors, with the PL4+ supporting up to 8 inputs (with expansion) and the PL4+D supporting up to 12 detectors (with remote cards). - Multiple relay outputs for alarms and faults, with the PL4+ offering 5 relays and the PL4+D managing up to 21 relays (with additional modules). - Visual and audible alarms to alert users to gas detection and system status. - Designed to be configurable for detecting various types of gases, including flammable, toxic, and oxygen levels.
PL4+
The PL4+ is a versatile gas control panel ideal for small systems. It supports 4 analog 4-20 mA inputs, expandable to 8 with an additional module. The unit includes 5 relays for alarms and faults, with an optional extension board offering more inputs and outputs. Housed in an IP65-rated ABS cabinet, the PL4+ features visual and audible alarms, and can accommodate a backup battery for uninterrupted operation.
PL4+
| Product Code | Description | PVP
| STPL4+ | 4-channel, triple threshold microprocessor-based gas control panel provided with LCD display, IP65 rated enclosure, 12V, 1.5A enclosed Power Supply and 5 output relays. It accepts 4-20 mA analog signals. Expandable to 8 inputs by using STPL4/ESP. | 1644.3
Accessories
| Product Code | Description | PVP
| STPL4/ESP | 4-channel extension card to expand the PL4+ control panel with 4 additional channels. It also provides 16 additional Open Collector Outputs. | 813.24
| STG/16REL | 16-relay card to be connected to MULTISCAN 8+ and STPL4/ESP to provide additional 16 relay outputs. * External box for STG/16REL not provided | 1028.7
| AL04 | 24 Vdc power supply for STPL4+ | 140.94
PL4+D
The PL4+D is designed for DIN-rail mounting and is perfect for small plants with up to 12 detectors. It can monitor 4 gas detectors directly and an additional 8 via remote input cards. The panel manages up to 21 output relays, ensuring comprehensive control and safety. Its backlit LCD display provides clear gas concentration readings and alarm set points, making it a reliable choice for diverse gas detection needs.
PL4+D | PL4+D
| Product Code | Description | PVP
| STPL4+D | 4 channel triple threshold DIN rail control panel. Possibility to add additional 8 inputs via STG/IN8S card or have the detectors connected to the RS485 bus. The panel comes complete with 1 USB port for PC connection. Power supply: 24 Vdc | 1644.3
Accessories
| Product Code | Description | PVP
| STG/IN4-OUT4-S1 | Remote module with 4 inputs 4-20mA and 4 output relays for MULTISCAN | 844.02
| STG/IN8S | High safety 8-input remote module with improved performance | 615.6
| STG/OUT16S | High safety 16-output remote module (Open collector type) with improved performance. | 615.6
| STG/8REL | Plug and play relay board to be connected to STG/OUT16 to provide 8 relay outputs | 516.78
| AL09 | 24 Vdc power supply for STPL4+D & STMTS/RTU | 165.24
| STG/IPBOX | IP65 ABS box for STPL4+D & STMTS/RTU | 268.92

## MULTISCAN 8+
Auto+Hide+Values
MULTISCAN 8+
The MULTISCAN 8+ is an advanced gas control unit, perfect for small plants with up to 16 or 24 detectors. It is highly flexible, capable of monitoring gas detectors connected via both 4-20 mA transmission and addressable bus. Each input can be configured for different types of gases, allowing detection of flammable and toxic compounds, as well as oxygen depletion and enrichment. The backlit LCD display provides precise readings of gas concentrations and alarm set points. The MULTISCAN 8+ ABS cabinet can house 2 x 12 VDC, 7Ah batteries to ensure proper operation during power failures. It is SIL1 approved and ATEX certified according to EN60079-29-1.
Multiscan 8+
| Product Code | Description | PVP
| STMTS/8+ | MULTISCAN 8+ gas control panel. 8 Analog 4-20 mA inputs and 6 relay outputs on-board the panel. 1 x RS485 type open bus line to connect additional 8 detectors through remote 8-input module or addressable gas detectors daisy chained to the RS485 bus.6 relay outputs and additional 16 relays via the dedicated STS/16REL card or via remote modules. The panel comes complete with 1 RS232 serial port and 1 USB port for PC connection. Housed in ABS box. | 3286.98
| STMTS/8+16 | MULTISCAN 8+16 gas control panel. 8 Analog 4-20 mA inputs and 6 relay outputs on-board the panel. 1 x RS485 type open bus line to connect additional 16 detectors through remote 8-input module or addressable gas detectors daisy chained to the RS485 bus. 6 relay outputs and additional 16 relays via the dedicated STS/16REL card or via remote modules. The panel comes complete with 1 RS232 serial port and 1 USB port for PC connection. Housed in ABS box. | 3900.96
Accessories
| Product Code | Description | PVP
| STG/IN4-OUT4-S1 | Remote module with 4 inputs 4-20mA and 4 output relays for MULTISCAN | 844.02
| STG/IN8S | High safety 8-input remote module with improved performance | 615.6
| STG/16REL | 16-relay card to be connected to MULTISCAN 8+ and STPL4/ESP to provide additional 16 relay outputs | 1028.7
| STG/OUT16S | High safety 16-output remote module (Open collector type) with improved performance. | 615.6
| STG/8REL | Plug and play relay board to be connected to STG/OUT16 to provide 8 relay outputs | 516.78
| STG/TCPIP-8+ | TCP/IP interface for LAN connection specific for MULTISCAN 8+ | 693.36
| STMTS/RTU | Din rail mounted remote unit for events visualization. To be RS485 connected to MULTISCAN++S1, PK & MULTISCAN 8+ serial lines, it allows the visualization and identification of events. RTU comes complete with an LCD display to show different information as alarm or fault, 4 analogue 4-20 mA inputs for detectors connection and 5 output relays freely programmable. | 2055.78
| AL09 | 24 Vdc power supply for STPL4+D & STMTS/RTU | 165.24
| STG/IPBOX | IP65 ABS box for STPL4+D & STMTS/RTU | 268.92
Model
RS-485 Bus Lines
Max. No. of Inputs
Max. No. of Outputs
STG/IN8S Modules
STG/OUT16S Modules
Multiscan++ 8+
1 open
8+8
6+16+16
1
1
Multiscan++ 8+16
1 open
8+16
6+16+32
2
2

## MULTISCAN++ S1
Auto+Hide+Values
MULTISCAN++ S1
The MULTISCAN++ S1 is a state-of-the-art gas control unit designed to monitor and control up to 256 gas detectors. It is extremely flexible and highly reliable, capable of overseeing detectors connected via both 4-20 mA transmission and addressable loops. The unit can be customized by connecting various field modules to the RS485 loops, adapting to specific application needs. Each remote input/output card connected to the loops features opto-isolation, safeguarding the system from electrical noise. The unit includes extensive self-testing capabilities to identify and pinpoint potential faults. The backlit graphic LCD display provides accurate readings of gas concentrations. Additionally, an RS232C serial port allows for PC connection to the CPU board for configuration and supervision. Event memory reports or real-time printouts can be generated by connecting a printer to the parallel printer port. The panels are available in either an IP65 ABS box or a 19” 6U rack. • SIL1 and ATEX according to EN60079-29-1. • Up to 4 open RS485 (EIA-485) or 2 closed galvanic isolated serial lines. • Up to 256 inputs (both 4-20 mA and loop daisy-chained) • Up to 512 O/C or relay outputs via remote modules and 8 relays on-board the control unit.
Multiscan++ S1
| Product Code | Description | PVP
| STMTS/S1-256B | MULTISCAN++S1 gas control panel. 4 x RS485 type open bus lines or 2 closed loops. 256 Analog 4-20 mA gas detectors connected through remote 8-input modules or addressable gas detectors daisy chained to the RS485 control panel’s loops and 8 x 4-20 mA inputs on-board the panel. 8 relay outputs and up to 512 O/C or relay outputs via remote modules. The panel comes complete with 1 parallel printer port and 1 RS232 serial port for PC connection. Housed in IP65 rated ABS box. | 6774.84
| STMTS/S1-256 | Rack version of STMTS/S1-256B. The panel comes housed in a 19” 6 U rack (24 Vdc power supply is required). | 6774.84
| STMTS/S1-128B | MULTISCAN++ S1 gas control panel. 4 x RS485 type open bus lines or 2 closed loops. 128 Analog 4-20 mA gas detectors connected through remote 8-input modules or addressable gas detectors daisy chained to the RS485 control panel’s loops and 8 x 4-20 mA inputs on-board the panel. 8 relay outputs and up to 256 O/C or relay outputs via remote modules. The panel comes complete with 1 parallel printer port and 1 RS232 serial port for PC connection. Housed in IP65 rated ABS box. | 6160.86
| STMTS/S1-128 | Rack version of STMTS/S1-128B. The panel comes housed in a 19” 6 U rack (24 Vdc power supply is required). | 6160.86
| STMTS/S1-64B | MULTISCAN++ S1 gas control panel. 2 x RS485 type open bus lines or 1 closed loop. 64 Analog 4-20 mA gas detectors connected through remote 8-input modules or addressable gas detectors daisy chained to the RS485 control panel’s loops and 8 x 4-20 mA inputs on-board the panel. 8 relay outputs and up to 128 O/C or relay outputs via remote modules. The panel comes complete with 1 parallel printer port and 1 RS232 serial port for PC connection. Housed in IP65 rated ABS box. | 5542.02
| STMTS/S1-64 | Rack version of STMTS/S1-64B. The panel comes housed in a 19” 6 U rack (24 Vdc power supply is required). | 5542.02
| STMTS/S1-32B | MULTISCAN++ S1 gas control panel. 2 x RS485 type open bus lines or 1 closed loop. 32 Analog 4-20 mA gas detectors connected through remote 8-input modules or addressable gas detectors daisy chained to the RS485 control panel’s loops and 8 x 4-20 mA inputs on-board the panel. 8 relay outputs and up to 64 O/C or relay outputs via remote modules. The panel comes complete with 1 RS232 serial port for PC connection. Housed in IP65 rated ABS box. | 5133.78
| STMTS/S1-32 | Rack version of STMTS/S1-32B. The panel comes housed in a 19” 6 U rack (24 Vdc power supply is required). | 5133.78
Accessories
| Product Code | Description | PVP
| STG/IN4-OUT4-S1 | Remote module with 4 inputs 4-20mA and 4 output relays for MULTISCAN++S1 | 844.02
| STG/IN8S | High safety 8-input remote module with improved performance | 615.6
| STG/OUT16S | High safety 16-output remote module (Open collector type) with improved performance. | 615.6
| STG/8REL | Plug and play relay board to be connected to STG/OUT16 to provide 8 relay outputs | 516.78
| STG/SER++ | Additional Serial board for MULTISCAN++S1-64 & 64B | 275.4
| STG/TCPIP-C | TCP/IP interface for LAN connection | 591.3
| AL11/276V5 | Power supply 27.6 Vdc 5A in Metal cabinet | 1057.86
| STG/BOX | IP55 Plastic box to house one STG/IN8-S or STG/OUT16-S card | 168.48
| STG/32-85 | RS232/RS485 Interface for PC connection on RS485 line of MULTISCAN++ | 289.98
| STMTS/RTU | Din rail mounted remote unit for events visualization. To be RS485 connected to MULTISCAN++S1, PK & MULTISCAN 8+ serial lines, it allows the visualization and identification of events. RTU comes complete with an LCD display to show different information as alarm or fault, 4 analogue 4-20 mA inputs for detectors connection and 5 output relays freely programmable. | 2055.78
| AL09 | 24 Vdc power supply for STPL4+D & STMTS/RTU | 165.24
| STG/IPBOX | IP65 ABS box for STPL4+D & STMTS/RTU | 268.92
Model
RS-485 Bus Lines
Max. No. of Inputs
Max. No. of Outputs
STG/IN8S Modules
STG/OUT16S Modules
Multiscan++ S1-256
4 open / 2 closed
8+256
8+512
32
32
Multiscan++ S1-128
4 open / 2 closed
8+128
8+256
16
16
Multiscan++ S1-64
2 open / 1 closed (+2 exp)
8+64
8+128
8
8
Multiscan++ S1-32
2 open / 1 closed
8+32
8+64
4
4

## MULTISCAN++ S2
Auto+Hide+Values
MULTISCAN++ S2
The MULTISCAN++ S2 is a cutting-edge gas control unit designed to monitor and manage up to 256 gas detectors. It is highly flexible and reliable, capable of overseeing detectors connected via both 4-20 mA transmission and addressable loops. The unit can be customized by connecting various field modules to the RS485 loops, adapting to specific application needs. Each remote input/output card connected to the loops features opto-isolation, safeguarding the system from electrical noise. The unit includes extensive self-testing capabilities to identify and pinpoint potential faults. The backlit graphic LCD display provides accurate readings of gas concentrations. Additionally, an RS232C serial port allows for PC connection to the CPU board for configuration and supervision. Event memory reports or real-time printouts can be generated by connecting a printer to the parallel printer port. The panels are available in either an IP65 ABS box or a 19” 6U rack. • ATEX EN6079-29-1 and SIL 2 certified (EN 50402) • Up to 2 closed RS485 (EIA-485) galvanic isolated loops • Up to 256 inputs (both 4-20 mA or addressable loop daisy-chained) • Up to 512 O/C outputs on remote modules (relay card optional) and 8 relays on the control unit’s board.
Multiscan++ S2
| Product Code | Description | PVP
| STMTS/S2-256B | MULTISCAN++ S2 Control Panel. 2 x RS485 type closed loops. 256 Analog 4-20 mA gas detectors connected through remote 8-input modules or addressable gas detectors daisy chained to the RS485 control panel’s loops and 8 x 4-20 mA inputs onboard the panel.8 relay outputs and up to 512 O/C or relay outputs via remote modules. The panel comes complete with 1 parallel printer port and 1 RS232 serial port for PC connection. Housed in IP65 rated ABS box | 7351.56
| STMTS/S2-256 | Rack version of STMTS/S2-256B. The panel comes housed in a 19” 6 U rack (24 V DC power supply is required). | 7351.56
| STMTS/S2-128B | MULTISCAN++ S2 control panel. 2 x RS485 type closed loops. 128 Analog 4-20 mA gas detectors connected through remote 8-input modules or addressable gas detectors daisy chained to the RS485 control panel’s loops and 8 x 4-20 mA inputs onboard the panel. 8 relay outputs and up to 256 O/C or relay outputs via remote modules. The panel comes complete with 1 parallel printer port and 1 RS232 serial port for PC connection. Housed in IP65 rated ABS box. | 6669.54
| STMTS/S2-128 | Rack version of STMTS/S2-128B. The panel comes housed in a 19” 6 U rack (24 Vdc power supply is required). | 4117
| STMTS/S2-64B | MULTISCAN++ S2 gas control panel. 2 x RS485 type closed loop. 64 Analog 4-20 mA gas detectors connected through remote 8-input modules or addressable gas detectors daisy chained to the RS485 control panel’s loops and 8 x 4-20 mA inputs onboard the panel. 8 relay outputs and up to 128 O/C or relay outputs via remote modules. The panel comes complete with 1 parallel printer port and 1 RS232 serial port for PC connection. Housed in IP65 rated ABS box. | 5963.22
| STMTS/S2-64 | Rack version of STMTS/S2-64B. The panel comes housed in a 19” 6 U rack (24 Vdc power supply is required). | 5963.22
| STMTS/S2-32B | MULTISCAN++ S2 gas control panel. 1 x RS485 type closed loop. 32 Analog 4-20 mA gas detectors connected through remote 8-input modules or addressable gas detectors daisy chained to the RS485 control panel’s loops and 8 x 4-20 mA inputs onboard the panel. 8 relay outputs and up to 64 O/C or relay outputs via remote modules. The panel comes complete with 1 RS232 serial port for PC connection. Housed in IP65 rated ABS box. | 5514.48
| STMTS/S2-32 | Rack version of STMTS/S2-32B. The panel comes housed in a 19” 6 U rack (24 Vdc power supply is required). | 5514.48
Accessories
| Product Code | Description | PVP
| STG/IN8-S2 | 8 input remote module card for MULTISCAN++S2 control panel. | 704.7
| STG/OUT16-S2 | 16 Open Collectors output remote module for MULTISCAN++S2 control panel. | 704.7
| STG/8REL | Plug & play relay board to be connected to STG/OUT16-S2 for 8 relay outputs | 516.78
| STG/TCPIP-C | TCP/IP interface for LAN connection | 591.3
| AL11/276V5 | Power supply 27.6 Vdc 5A in Metal cabinet | 1057.86
| STG/BOX | IP55 Plastic box to house one STG/IN8-S2 or STG/OUT16-S2 cards | 168.48
| STG/32-85 | RS232/RS485 Interface for PC connection on RS485 line of MULTISCAN++ | 289.98
Model
RS-485 Bus Lines
Max. No. of Inputs
Max. No. of Outputs
STG/IN8S Modules
STG/OUT16S Modules
Multiscan++ S1-256
4 open / 2 closed
8+256
8+512
32
32
Multiscan++ S1-128
4 open / 2 closed
8+128
8+256
16
16
Multiscan++ S1-64
2 open / 1 closed (+2 exp)
8+64
8+128
8
8
Multiscan++ S1-32
2 open / 1 closed
8+32
8+64
4
4

## SECURNET++ & SCADA LITE
Auto+Hide+Values
SECURNET++ & SCADA LITE Supervising Software
SECURNET++ The SECURNET++ system is a supervising software that allows to monitor the status of a gas detection system equipped with SENSITRON gas detection panels, model MULTISCAN++S1, MULTISCAN++S2, GALILEO SMS and GALILEO 8/32. SECURNET++ allows to: Manage the gas detection system directly or connected on a Lan with TCP/IP protocol by optional module. Display the status and concentration of the gas sensors on graphic maps that are hierarchically organised. Display the status of the digital output channels (siren, solenoid valves etc.) on graphic maps that are hierarchically organised. Display the gas concentration measured by the sensors in numerical form (display), analogical bar-graph or trend (diagram) as a function of time. Check the historical log event of the system Run under the Windows XP, Vista, WIN7, 8, 8.1, 10 Use of ModBus protocol to communicate with the panels Multi-language support (Italian and English languages are pre-set). Configuration and import capabilities for the graphical maps (screens) High level of configuration potential and flexibility Automatic activation of the map where the last alarm occurred Easy and intuitive interaction with the user. Password protection included SCADA LITE The SCADA LITE application installed in the Windows environment allows you to monitor 1 control unit to receive, through the Modbus protocol, the alarm and the fault information, highlighting them on an event list with real-time update. SCADA LITE allows you to manage a remote control system (RS232 serial) or on a local TCP / IP network using an optional module. The SCADA LITE software runs on a mini-PC with touch screen. It allows the creation of a remote terminal for the repetition and management of a gas detection system with a MULTISCAN ++ control unit.
Securnet++
| Product Code | Description | PVP
| ST.SW/S++ | SECURNET++ Map graphic software for the remote monitoring and management of one gas detection plant. Running under Windows WIN7, 8, 8.1, 10. Modbus protocol for the communication with the control panels. Trend and bar-graph performance of the detectors. | 11534.4
| ST.SW/S++1 | SECURNET++ License for two control units | 12879
| ST.SW/S++2 | SECURNET++ License for three control units | 14160.42
Scada Lite
| Product Code | Description | PVP
| ST.SW/SL | SCADA LITE event reception software for the management of 1 gas detection system with Multiscan ++ control unit. Operated with Windows XP, Vista, 7, 8, 8.1, 10. It is supplied on an USB flash drive with protection. | Free of charge

## MULTISCAN++ MED
Auto+Hide+Values
MULTISCAN++ MED Centrale per applicazioni Marine
The MULTISCAN++ MED system, equipped with a powerful microprocessor, can manage up to 64 gas detectors and is tailored to meet the needs of the marine industry. It is certified by Lloyd’s Register under MED/3.54 Regulation and also holds ATEX and SIL2 certifications. This highly flexible system includes 8 preset relays and can support an additional 128 open collector outputs through remote STG/OUT16-S2 output cards. Each device connected to the loops is galvanically isolated to protect the system from electrical noise. The system incorporates extensive self-testing features to detect and pinpoint faults. The backlit graphic LCD display provides accurate gas concentration readings. An RS232C serial port allows for PC connection to the CPU board for configuration and supervision. Event memory reports or real-time printouts can be generated by connecting a printer to the parallel printer port. The panels are housed in a durable metal cabinet. • Marine, ATEX and SIL 2 certified (EN 50402) • Up to 2 closed RS485 (EIA-485) galvanic isolated loops • Up to 64 inputs (standard configuration) and up to 256 inputs on request • Up to 128 O/C outputs (standard configuration) and up to 256 outputs on request.
Multiscan++ MED
| Product Code | Description | PVP
| STMTS/S2-64MED | MULTISCAN++ MED SIL2 and MED approved gas control panel. 1 x RS485 type closed loop. 8+8 4-20 mA gas detectors directly connected to the panel and up to total 64 detectors connectable through analogue inputs Modules or Addressable gas detectors on the RS485 control panel’s loops. 8+8 relay output and up to 128 by O/C Outputs on remote modules. Dedicated output for parallel printer. 1 RS232 serial port for PC connection. Housed in stout metal cabinet. | 18058.14
Accessories
| Product Code | Description | PVP
| STG/IN8-S2MED | 8 input remote module card for MULTISCAN++MED control panel. | 822.96
| STG/OUT16-S2MED | 16 Open Collectors output remote module for MULTISCAN++MED control panel. | 822.96
| STG/8REL | Plug and play relay board to be connected to STG/OUT16S to provide 8 relay outputs | 516.78
Model
RS-485 Bus Lines
STG/IN8-S2MED
STG/OUT16-S2MED
MULTISCAN++MED
1 closed
8+8+64
8+128
8
8
Max. No. of Inputs
Max. No. of Outputs

## COVER PK SYSTEM
#VALUE!

## SMART P
Auto+Hide+Values
SMART P Gas Detectors for Parking Applications
The SMART P gas detectors are designed to meet the new European Standard for gas detection in car parks and are certified to comply with EN 50545-1 when connected to MULTISCAN++ PK panels. These detectors offer reliable gas detection with standard RS-485 communication, zero tracking for stability, continuous system self-diagnosis, and a 24 VDC power supply. Features: - Certified to EN 50545-1 when connected to MULTISCAN++ PK panels - Standard RS485 - Maintains zero point stability - Ongoing Self-Diagnosis system checks - 24 Vdc Power Supply - Must be connected to MULTISCAN++ PK Controller
Smart P-1
Smart P-1 is a single sensor gas detector for CO (0-300 ppm) or NO2 (0-30 ppm) with RS485 communication, designed for car park gas detection.
| Product Code | Description | PVP | Sensor Head |  | Gas Cylinder
| S2398CO | Smart P1, Carbon Monoxide CO EC 0-300 ppm, RS-485 | 510.3 | K2398CO |  | SIB56
| S2399ND | Smart P1, Nitrogen Dioxide NO2 0-30 ppm, RS-485 | 510.3 | K2399ND |  | RIB116-34
Smart P-2
Smart P-2 is a dual sensor gas detector capable of detecting CO & NO2 or CO & Petrol vapours, featuring RS485 communication and remote sensor head for comprehensive car park gas detection
| Product Code | Description | PVP | Sensor Head |  | Gas Cylinder
| S2400CO-ND | Smart P2, Carbon Monoxide CO EC 0-300 ppm & Nitrogen Dioxide NO2 EC 0-30 ppm, RS-485 | 845.64 | K2398CO | K2399ND | SIB56 + RIB116-34
| S2400CO-VB | Smart P2, Carbon Monoxide CO EC 0-300 ppm & Petrol Vapour Cat 0-100% LEL (Supplied with remote sensor head complete with 130cm cable) | 886.14 | K2398CO | K2400VB | SIB56 + SIB06
Accessories
| Product Code | Description | PVP
| STS/CKD+ | Enhanced calibration keypad with 7-digit display and SD card slot. Compatible with Smart series for operations like zero/span calibration and monitor 4-20 mA signals. | 515.16

## MULTISCAN++ PK
Auto+Hide+Values
MULTISCAN++ PK
The gas detection system for car parks includes MULTISCAN++ PK control panels and SMART-P gas detectors, specifically designed to meet the European norm EN 50545-1. It can manage and monitor up to 256 gas detectors and can be customized to meet various field requirements. Each gas detector can have three alarm thresholds. The panels allow alarms to be time-averaged over a period adjustable from 5 to 60 minutes for thresholds 1 and 2, while the third threshold triggers an alarm within 1 to 5 minutes. The panels feature 8 analog 4-20 mA inputs and 8 relays. The system architecture supports connection to remote 16-output cards (STG/OUT16-S) and 8-relay boards. Each remote input/output card connected to the loops features opto-isolation, safeguarding the system from electrical noise. The system includes extensive self-testing facilities to detect and localize faults. The backlit graphic LCD display provides precise gas concentration readings. An RS232C serial port allows for PC connection to the CPU board for configuration and supervision. Event memory reports or real-time printouts can be generated by connecting a printer to the parallel printer port. The panels are available in either an IP65 ABS box or a 19” 6U rack. • Designed according to the new European Standard for gas detection in car parks. • Up to 4 open RS485 (EIA-485) or 2 closed galvanic isolated serial lines • Up to 256 inputs (both 4-20 mA or addressable loop, daisy-chained) • Up to 512 O/C or relay outputs via remote modules and 8 relays on-board the control unit.
Multiscan++ PK
| Product Code | Description | PVP
| STMTS/PK-256B | MULTISCAN++ PK gas control panel. 4 x RS485 open bus lines or two closed loops. 256 Analog 4-20 mA gas detectors connected through remote 8-input modules or addressable gas detectors daisy chained to the RS485 control panel’s loops and 8 x 4-20 mA inputs on-board the panel.8 relay outputs and up to 512 O/C or relay outputs via remote modules. The panel comes complete with 1 parallel printer port and 1 RS232 serial port for PC connection. Housed in IP65 rated ABS box | 6774.84
| STMTS/PK-256 | Rack version of STMTS/PK-256B. The panel comes housed in a 19” 6 U rack (24 Vdc power supply is required). | 6774.84
| STMTS/PK-128B | MULTISCAN++ PK gas control panel. 4 x RS485 type open bus lines or two closed loops. 128 Analog 4-20 mA gas detectors connected through remote 8-input modules or addressable gas detectors daisy chained to the RS485 control panel’s loops and 8 x 4-20 mA inputs on-board the panel. 8 relay outputs and up to 256 O/C or relay outputs via remote modules. The panel comes complete with 1 parallel printer port and 1 RS232 serial port for PC connection. Housed in IP65 rated ABS box | 6160.86
| STMTS/PK-128 | Rack version of STMTS/PK-128B. The panel comes housed in a 19” 6 U rack (24 Vdc power supply is required). | 6160.86
| STMTS/PK-64B | MULTISCAN++ PK gas control panel. 2 x RS485 type open bus lines or one closed loop. 64 Analog 4-20 mA gas detectors connected through remote 8-input modules or addressable gas detectors daisy chained to the RS485 control panel’s loops and 8 x 4-20 mA inputs on-board the panel. 8 relay outputs and up to 256 O/C or relay outputs via remote modules. The panel comes complete with 1 parallel printer port and 1 RS232 serial port for PC connection. Housed in IP65 rated ABS box | 5542.02
| STMTS/PK-64 | Rack version of STMTS/PK-64B. The panel comes housed in a 19” 6 U rack (24 Vdc power supply is required). | 5542.02
| STMTS/PK-32B | MULTISCAN++ PK gas control panel. 2 x RS485 type open bus lines or one closed loop. 32 Analog 4-20 mA gas detectors connected through remote 8-input modules or addressable gas detectors daisy chained to the RS485 control panel’s loops and 8 x 4-20 mA inputs on-board the panel. 8 relay outputs and up to 256 O/C or relay outputs via remote modules. The panel comes complete with 1 parallel printer port and 1 RS232 serial port for PC connection. Housed in IP65 rated ABS box | 5133.78
| STMTS/PK-32 | Rack version of STMTS/PK-32B. The panel comes housed in a 19” 6 U rack (24 Vdc power supply is required). | 5133.78
| STMTS/PK-8D | MULTISCAN++ PK gas control panel. 1 x RS485 type open bus line. 8 addressable gas detectors daisy chained to the RS485 control panel’s loops and 4 x 4-20 mA inputs on-board the panel. 5 relay outputs and additional 16 via remote STG/OUT16S card. The panel comes complete with 1 USB port for PC connection. Housed in plastic box for din-rail mounting | 2055.78
Accessories
| Product Code | Description | PVP
| STG/IN4-OUT4-PK | Remote module with 4 inputs 4-20mA and 4 output relays for MULTISCAN++PK | 844.02
| STG/IN8-PK | 8 input remote module for Multiscan++ PK control panel. | 615.6
| STG/OUT16S | High safety 16-output remote module (Open collector type) with improved performance. Max 8 modules per line, for a total of 16 modules connectable to the system. For STMTS/PK-64B and STMTS/PK-32B max 8 modules connectable to the system | 615.6
| STG/8REL | Plug and play relay board to be connected to STG/OUT16 to provide 8 relay outputs | 516.78
| STG/SER++ | Additional Serial board for MULTISCAN++PK-64 & 64B | 275.4
| STG/TCPIP-C | TCP/IP interface for LAN connection | 591.3
| AL11/276V5 | Power supply 27.6 Vdc 5A in Metal cabinet | 1057.86
| STG/BOX | IP55 Plastic box to house one STG/IN8-PK or STG/OUT16S card | 168.48
| STG/32-85 | RS232/RS485 Interface for PC connection on RS485 line of MULTISCAN++ | 289.98
| STMTS/RTU | Din rail mounted remote unit for events visualization. To be RS485 connected to MULTISCAN++S1, PK & MULTISCAN 8+ serial lines, it allows the visualization and identification of events. RTU comes complete with an LCD display to show different information as alarm or fault, 4 analogue 4-20 mA inputs for detectors connection and 5 output relays freely programmable. | 2055.78
| AL09 | 24 Vdc power supply for STPL4+D & STMTS/RTU | 165.24
| STG/IPBOX | IP65 ABS box for STPL4+D & STMTS/RTU | 268.92

## Accessories for SMART Serie
Auto+Hide+Values
Accessories for SMART Detector Series
Installation Accessories
| Product Code | Description | PVP
| ZMTEST2 | Permanent rain shield/test adapter suitable for sensor head type 2, (like S2096__ and S2097__) | 69.66
| ZMTEST3 | Permanent rain shield/test adapter suitable for sensor head type 3 | 90.72
| SL517 | Stainless steel collector and weather protection cone for Explosion proof gas detector (suitable for sensor head type 3). | 45.36
| SL523 | Stainless steel collector and weather protection cone for Explosion proof gas detector suitable for sensor head type 2 (S2096__ and S2097__) | 66.42
| SL647 | Stainless steel weather protection roof, to protect gas detectors with vertical mounted sensor head and Stainless steel 2" mounting bracket included. Stainless steel AISI 316L. Suitable for SMART S and SMART3G | 925.02
| SL673 | Stainless steel 2" mounting bracket | 194.4
| STGAD.COND | Duct mounting adaptor | 408.24
Test Kits
| Product Code | Description | PVP
| VTX+XA | Calibration gas case including one flow indicator, hose and adaptor. (Calibration gas cylinders not included). | 432.54
| SIBXXX | Disposable 12 litre gas cylinders filled with either flammable or toxic non reactive gas mixtures. At the order please specify the gas type. Dangerous goods: ADR/RID UN 2037- UN 1950, Class 2, Packing instructions 203 | See Gas Cylinders page
| RIBXXX-34 | Disposable 34 litre gas cylinders filled with reactive gas mixtures. At the order please specify the gas type. Dangerous goods: ADR/RID UN 1956, Class 2. | See Gas Cylinders page
| PCVALVOLA | Flow-meter valve to be mounted on 12 liter gas cylinders SIBXXX to set the gas flow | 155.52
| MINIFLOW-34-110-4010 | Flow-meter valve to be mounted on 34 liter gas cylinders RIBXXX-34 to set the gas flow | 309.42
| ZMCAP2 | Adapter for detectors with sensor head type 2 in order to fit the duct mounting adapter accessory STGAD.COND | 111.78
| ZMCAP/123 | Aluminum calibration adapter for all of Sensitron gas detectors. It allows the right quantity of gas inlet to flow to the detector. It comes complete with adapters to fit on the detectors' heads. | 155.52
Stainless Steel Collector
Calibration Case (VTX+XA)
|  | Calibration Adapter (ZMCAP/123) |  | Flow Meter Valve (PCVALVOLA)

## Spare Parts
Auto+Hide+Values
Spare Sparts
| Product Code | Decription | PVP
| K1450CO | CARBON MONOXIDE / MONOSSIDO DI CARBONIO (CO) , 0-300 ppm | 289.98
| K1451ND | NITROGEN DIOXIDE / BIOSSIDO DI AZOTO (NO2), 0-20 ppm | 387.18
| K1455GP | LPG / GPL, 4-20 mA, 0-100% LFL | 307.8
| K1455ME | METHANE / METANO 4-20 mA, 0-100% LFL | 307.8
| K1455PR | PROPANE / PROPANO, 4-20 Ma, 0-100% LFL | 307.8
| K1455VB | PETROL VAPOURS 7 VAPORI BENZINA, 0-100% LFL , | 307.8
| K2398CO | CARBON MONOXIDE / MONOSSIDO DI CARBONIO (CO) , RS485. Range 0-300 ppm | 314.28
| K2399ND | NITROGEN DIOXIDE / BIOSSIDO DI AZOTO (NO2), RS485. Range 0-30 ppm. | 396.9
| K2400VB | PETROL VAPOURS / VAPORI BENZINA, RS485. Range 0-100% LFL | 497.34
| K2401CO2 | CARBON DIOXIDE / ANIDRIDE CARBONICA (CO2) , 0-5000 ppm | 1099.98
| K3004SD | SULPHUR DIOXIDE / ANIDRIDE SOLFOROSA (SO2) , 0-20 ppm | 1579.5
| K3008ND | NITROGEN DIOXIDE / BIOSSIDO DI AZOTO (NO2) , 0-20 ppm | 1579.5
| K3014CL | CHLORINE / CLORO (CL2) , 0-20 ppm | 1579.5
| K3017R410A | R410A , 0-2000 ppm | 2156.22
| K3229AM | AMMONIA / AMMONIACA, 0-5000 ppm | 1077.3
| K3425R32 | R32 DIFLUOROMETHANE detector , 0-100% LFL | 2156.22
| K3429R1234YF | R1234yf , 0-100% LFL | 2156.22
| K3433R1234ZE | R1234ze , 0-100% LFL | 2156.22
| K3593H2 | HYDROGEN / IDROGENO, 0-2000 ppm | 1334.88
| K3593H2 | HYDROGEN / IDROGENO, 0-2000 ppm | 1334.88
| K3596R134A | R134a , 0-2000 ppm | 2156.22
| K3599SF6 | SF6 , 0-2000 ppm | 2156.22
| K3656R125 | R125 , 0-2000 ppm | 2196.72
| K3657R1234YF | R1234yf , 0-2000 ppm | 2156.22
| K3658R404A | R404A , 0-2000 ppm | 2156.22
| K3659R407A | R407A , 0-2000 ppm | 2156.22
| K3660R507A | R507A , 0-2000 ppm | 2156.22
| K3913CO | CARBON MONOXIDE (CO) detector, 0-500 ppm,4-20 mA, electrochem. cell | 753.3
| K3913CO | CARBON MONOXIDE (CO) detector, 0-500 ppm,4-20 mA, electrochem. cell | 753.3
| K3914ME | METHANE detector, 0-100% LFL, 4-20 mA, catalytic sensor | 575.1
| K3914ME | METHANE detector, 0-100% LFL, 4-20 mA, catalytic sensor | 575.1
| K3915O2 | OXYGEN (O2) detector, 0-25% vol, 4-20 mA, electrochemical cell | 753.3
| K3915O2 | OXYGEN (O2) detector, 0-25% vol, 4-20 mA, electrochemical cell | 753.3
| K3916ME | METHANE detector, 0-100% LFL, 4-20 mA, infrared sensor | 1294.38
| K3916ME | METHANE detector, 0-100% LFL, 4-20 mA, infrared sensor | 1294.38
| K3920R407F | R407F , 0-2000 ppm | 2156.22
| K3923R449A | R449A , 0-2000 ppm | 2156.22
| K3923R449A | R449 , 0-2000 ppm | 2156.22
| K3944R32 | R32 , 0-2000 ppm | 2156.22
| K3948R1234ZE | R1234ze , 0-2000 ppm | 2156.22
| K3952R417A | R417A , 0-2000 ppm | 2156.22
| K3956R448 | R448 , 0-2000 ppm | 2156.22
| K4433R454B | R454B , 0-100% LFL | 2156.22
| K4547H2 | HYDROGEN / IDROGENO, 0-3000 ppm | 1334.88
| K4547H2 | HYDROGEN / IDROGENO, 0-3000 ppm | 1334.88
| K4590BU | BUTANE /- BUTANO , 0-100% LFL | 445.5
| K4590GP | LPG / GPL , 0-100% LFL | 445.5
| K4590ME | METHANE / METANO , 0-100% LFL | 445.5
| K4590PR | PROPANE / PROPANO , 0-100% LFL | 445.5
| K4590VB | PETROL VAPOURS / VAPORI BENZINA , 0-100% LFL | 424.875
| K4591ME | METHANE / METANO , 0-100% LFL | 886.14
| K4592BU | BUTANE / BUTANO , 0-100% LFL | 886.14
| K4592DE | R152a , 0-100% LFL | 886.14
| K4592PR | PROPANE / PROPANO , 0-100% LFL | 886.14
| K4593O2 | OXYGEN / OSSIGENO , 0-25% Vol. | 622.08
| K4595CO | CARBON MONOXIDE / MONOSSIDO DI CARBONIO (CO) , 0-500 ppm | 622.08
| K4597AM | AMMONIA / AMMONIACA (NH3), 0-100 ppm | 871.56
| K4598AM | AMMONIA / AMMONIACA (NH3), 0-200 ppm | 871.56
| K4599AM | AMMONIA / AMMONIACA (NH3), 0-1000 ppm | 871.56
| K4601HS | HYDROGENE SULPHIDE / IDROGENO SOLFORATO (H2S) , 0-50 ppm | 604.26
| K4606CO2 | CARBON DIOXIDE / ANIDRIDE CARBONICA (CO2), 0-2% Vol. | 920.16
| K4607CO2 | CARBON DIOXIDE / ANIDRIDE CARBONICA (CO2), 0-5% Vol. | 920.16
| K4608CO2 | CARBON DIOXIDE / ANIDRIDE CARBONICA (CO2), 0-30% Vol. | 920.16
| K4610CO2 | CARBON DIOXIDE / ANIDRIDE CARBONICA (CO2), 0-5000ppm | 920.16
| K4611R32 | R32 , 0-2000 ppm | 1449.9
| K4612R134A | R134A , 0-2000 ppm | 1449.9
| K4613SF6 | SF6 , 0-2000 ppm | 1449.9
| K4614R125 | R125 , 0-2000 ppm | 1449.9
| K4615R1234YF | R1234yf , 0-2000 ppm | 1449.9
| K4616R1234ZE | R1234ze , 0-2000 ppm | 1449.9
| K4617R404A | R404A , 0-2000 ppm | 1449.9
| K4618R407A | R407A , 0-2000 ppm | 1449.9
| K4619R407F | R407F , 0-2000 ppm | 1449.9
| K4620R449A | R449A , 0-2000 ppm | 1449.9
| K4621R507A | R507A , 0-2000 ppm | 1449.9
| K4622R410A | R410A , 0-2000 ppm | 1449.9
| K4623R417A | R417A , 0-2000 ppm | 1449.9
| K4624R448 | R448 , 0-2000 ppm | 1449.9
| K4626R32 | R32 DIFLUOROMETHANE , 0-100% LFL | 1475.82
| K4627R1234YF | R1234yf , 0-100% LFL | 1475.82
| K4628R1234ZE | R1234ze , 0-100% LFL | 1475.82
| K4629R454B | R454B , 0-100% LFL | 1475.82
| K4630SD | SULPHUR DIOXIDE / ANIDRIDE SOLFOROSA (SO2) , 0-20 ppm | 994.68
| K4632ND | NITROGEN DIOXIDE / BIOSSIDO DI AZOTO (NO2) , 0-20 ppm | 994.68
| K4634CL | CHLORINE / CLORO (Cl2) , 0-20 ppm | 994.68
| K4638R32 | R32 , 0-5000 ppm | 871.56
| K4639R134A | R134a , 0-5000 ppm | 871.56
| K4640R404A | R404A , 0-5000 ppm | 871.56
| K4641R407C | R407C , 0-5000 ppm | 871.56
| K4642R410A | R410A , 0-5000 ppm | 871.56
| K4643R1234YF | R1234yf , 0-5000 ppm | 871.56
| K4644R1234ZE | R1234ze , 0-5000 ppm | 871.56
| K4827R454C | R454C , 0-100% LFL | 2156.22
| K4827R454C | R454C , 4-20mA, 0-100% LFL | 2156.22
| K4857R1233ZD | R1233Zd , 0-5000 ppm | 871.56
| K4949H2 | HYDROGEN / IDROGENO , 0-100%LFL (MPS) | 1508.22
| K4949ME | METHANE / METANO , 0-100%LFL (MPS) | 1508.22
| K5116R32 | SMART 3 Blizzard R32, 0 - 100% LFL | 703.08
| K5117R454B | SMART 3 Blizzard R454B, 0 - 100% LFL | 703.08
| K5118R410A | SMART 3 Blizzard R410A, 0 - 10.000 ppm | 703.08
| K5226R1234YF | SMART 3 Blizzard R1234YF, 0 - 5.000 ppm | 810
| K5227R1234ZE | SMART 3 Blizzard R1234ZE, 0 - 5.000 ppm | 810
| K5228R134A | SMART 3 Blizzard R134A, 0 - 5.000 ppm | 810
| K5229R404A | SMART 3 Blizzard R404A, 0 - 5.000 ppm | 810
| Kx097AA | ACETIC ACID / ACIDO ACETICO , 0-100% LFL | 515.16
| Kx097AB | BUTYL ACETATE / ACETATO DI BUTILE , 0-100% LFL | 515.16
| Kx097AE | ETHYL ACETATE / ACETATO DI ETILE , 0-100% LFL | 515.16
| Kx097AT | ACETONE , 0-100% LFL | 515.16
| Kx097BT | BUTYL ALCOHOL / ALCOL BUTILICO , 0-100% LFL | 515.16
| Kx097BU | BUTANE / BUTANO , 0-100% LFL | 515.16
| Kx097CE | CYCLOHEXANE / CICLOESANO , 0-100% LFL | 515.16
| Kx097CP | CYCLOPENTANE / CICLOPENTANO , 0-100% LFL | 515.16
| Kx097EE | ETHYL ETHER / ETERE ETILICO , 0-100% LFL | 515.16
| Kx097EL | ETHYLENE / ETILENE , 0-100% LFL | 515.16
| Kx097EN | ETHANE / ETANO , 0-100% LFL | 515.16
| Kx097EP | HEPTANE / EPTANO , 0-100% LFL | 515.16
| Kx097ES | HEXANE / ESANO , 0-100% LFL | 515.16
| Kx097ET | ETHYL ALCOHOL / ALCOL ETILICO , 0-100% LFL | 515.16
| Kx097GP | LPG / GPL , 0-100% LFL | 515.16
| Kx097H2 | HYDROGEN / IDROGENO , 0-100% LFL | 515.16
| Kx097IB | ISO BUTANE / ISO BUTANO , 0-100% LFL | 515.16
| Kx097IN | ISO PENTANE / ISO PENTANO , 0-100% LFL | 515.16
| Kx097IP | ISO PROPYL ALCOHOL / ALCOL ISO PROPILICO , 0-100% LFL | 515.16
| Kx097IT | ISO BUTYL ALCOHOL / ALCOL ISO BUTILICO , 0-100% LFL | 515.16
| Kx097ME | METHANE / METANO , 0-100% LFL | 515.16
| Kx097MK | METHYL ETHYL KETONE / METIL ETIL CHETONE , 0-100% LFL | 515.16
| Kx097MT | METHYL ALCOHOL / ALCOL METILICO , 0-100% LFL | 515.16
| Kx097NN | NONANE / NONANO , 0-100% LFL | 515.16
| Kx097PE | PENTANE / PENTANO , 0-100% LFL | 515.16
| Kx097PN | PROPYL ALCOHOL / ALCOL PROPILICO , 0-100% LFL | 515.16
| Kx097PP | PROPYLENE / PROPILENE , 0-100% LFL | 515.16
| Kx097PR | PROPANE / PROPANO , 0-100% LFL | 515.16
| Kx097TM | TRI-METHYLBENZENE / TRIMETILBENZENE , 0-100% LFL | 515.16
| Kx097TO | TOLUENE , 0-100% LFL | 515.16
| Kx097VB | PETROL VAPOURS / VAPORI BENZINA , 0-100% LFL | 515.16
| Kx097XI | XYLENE / XILENE , 0-100% LFL | 515.16
| Kx129O2 | OXYGEN / OSSIGENO, 0-30% Vol. | 753.3
| Kx131CO | CARBON MONOXIDE / MONOSSIDO DI CARBONIO, 0-500 ppm | 753.3
| Kx131CO | CARBON MONOXIDE / MONOSSIDO DI CARBONIO, 0-500 ppm | 753.3
| Kx133AM | AMMONIA / AMMONIACA, 0-200 ppm | 1077.3
| Kx135AM | AMMONIA / AMMONIACA, 0-1000 ppm | 1077.3
| Kx137HS | HYDROGEN SULPHIDE / IDROGENO SOLFORATO (H2S), 0-50 ppm | 863.46
| Kx145NO | NITRIC OXIDE / MONOSSIDO DI AZOTO (NO), 0-100 ppm | 1334.88
| Kx151AC | ACETYLENE / ACETILENE , 0-100% LFL | 575.1
| Kx151AC | ACETYLENE / ACETILENE , 0-100% LFL | 575.1
| Kx154P8 | JP8 , 0-100% LFL detector | 607.5
| Kx154P8 | JP8 (JET PROPELLER) , 0-100% LFL | 607.5
| Kx155AA | ACETIC ACID / ACIDO ACETICO , 0-100% LFL | 575.1
| Kx155AE | ETHYL ACETATE / ACETATO DI ETILE , 0-100% LFL | 575.1
| Kx155AM | AMMONIA / AMMONIACA , 0-100% LFL | 575.1
| Kx155AM | AMMONIA / AMMONIACA , 0-100% LFL | 575.1
| Kx155AT | ACETONE , 0-100% LFL | 575.1
| Kx155AV | VINYL ACETATE / ACETATO DI VINILE , 0-100% LFL | 575.1
| Kx155BU | BUTANE / BUTANO , 0-100% LFL | 575.1
| Kx155BU | PROPANE / PROPANO , 0-100% LFL | 575.1
| Kx155BZ | BENZENE , 0-100% LFL | 575.1
| Kx155CE | CYCLOHEXANE / CICLOESANO , 0-100% LFL | 575.1
| Kx155CP | CYCLOPENTANE / CICLOPENTANO , 0-100% LFL | 575.1
| Kx155EE | ETHYL ETHER / ETERE ETILICO , 0-100% LFL | 575.1
| Kx155EL | ETHYLENE / ETILENE , 0-100% LFL | 575.1
| Kx155EN | ETHANE / ETANO , 0-100% LFL | 575.1
| Kx155EP | HEPTANE / EPTANO , 0-100% LFL | 575.1
| Kx155ES | HEXANE / ESANO , 0-100% LFL | 575.1
| Kx155ET | ETHYL ALCOHOL / ALCOL ETILICO , 0-100% LFL | 575.1
| Kx155GP | LPG / GPL , 0-100% LFL | 575.1
| Kx155H2 | HYDROGEN / IDROGENO , 0-100% LFL | 575.1
| Kx155IB | ISO BUTANE / ISO BUTANO , 0-100% LFL | 575.1
| Kx155IN | ISO PENTANE / ISO PENTANO , 0-100% LFL | 575.1
| Kx155IP | ISO PROPYL ALCOHOL / ALCOL ISO PROPILICO , 0-100% LFL | 575.1
| Kx155IT | ISO BUTYL ALCOHOL / ALCOL ISO BUTILICO , 0-100% LFL | 575.1
| Kx155ME | METHANE / METANO, 0-100% LFL | 575.1
| Kx155ME | METHANE / METANO, 0-100% LFL | 575.1
| Kx155MK | METHYL ETHYL KETONE 7 METIL ETIL CHETONE , 0-100% LFL | 575.1
| Kx155MT | METHYL ALCOHOL / ALCOL METILICO , 0-100% LFL | 575.1
| Kx155NN | NONANE / NONANO , 0-100% LFL | 575.1
| Kx155OE | ETHYLENE OXIDE / OSSIDO DI ETILENE , 0-100% LFL | 575.1
| Kx155PE | PENTANE / PENTANO , 0-100% LFL | 575.1
| Kx155PN | PROPYL ALCOHOL / ALCOL PROPILICO , 0-100% LFL | 575.1
| Kx155PP | PROPYLENE / PROPILENE , 0-100% LFL | 575.1
| Kx155PR | PROPANE / PROPANO , 0-100% LFL | 575.1
| Kx155PR | BUTANE / BUTANO , 0-100% LFL | 575.1
| Kx155ST | STYRENE / STIRENE , 0-100% LFL | 575.1
| Kx155ST | STYRENE / STIRENE , 0-100% LFL | 575.1
| Kx155TM | TRI-METHYLBENZENE / TRIMETILBENZENE , 0-100% LFL | 575.1
| Kx155TO | TOLUENE , 0-100% LFL | 575.1
| Kx155VB | PETROL VAPOURS / VAPORI BENZINA , 0-100% LFL | 575.1
| Kx155XI | XYLENE / XILENE , 0-100% LFL | 575.1
| Kx225AM | AMMONIA / AMMONIACA, 0-100 ppm | 1077.3
| Kx444CO2 | CARBON DIOXIDE / ANIDRIDE CARBONICA (CO2), 0-5.000ppm Vol. | 1294.38
| Kx444CO2 | CARBON DIOXIDE / ANIDRIDE CARBONICA (CO2) , 0-5.000 ppm | 1294.38
| Kx444CO2 | CARBON DIOXIDE / ANIDRIDE CARBONICA (CO2), 0-5000 ppm | 1294.38
| Kx445CO2 | CARBON DIOXIDE / ANIDRIDE CARBONICA (CO2), 0-30% Vol. | 1294.38
| Kx586CO | CARBON MONOXIDE / MONOSSIDO DI CARBONIO, 0-1000 ppm | 753.3
| Kx624ME | METHANE / METANO, 0-100% LFL | 1294.38
| Kx624ME | METHANE / METANO, 0-100% LFL | 1294.38
| Kx630CO2 | CARBCARBON DIOXIDE / ANIDRIDE CARBONICA (CO2), 0-100% Vol. | 1294.38
| Kx641O2 | OXYGEN / OSSIGENO, 0-25% vol | 753.3
| Kx641O2 | OXYGEN / OSSIGENO, 0-25% Vol. | 753.3
| Kx651CO2 | CARBON DIOXIDE / ANIDRIDE CARBONICA (CO2), 0-2% Vol. | 1294.38
| Kx652CO2 | CARBON DIOXIDE / ANIDRIDE CARBONICA (CO2), 0-5% Vol. | 1294.38
| Kx654AE | ETHYL ACETATE / ACETATO DI ETILE , 0-100% LFL | 1294.38
| Kx654AT | ACETONE , 0-100% LFL | 1294.38
| Kx654BN | BUTADIENE , 0-100% LFL | 1294.38
| Kx654BU | BUTANE / BUTANO , 0-100% LFL | 1294.38
| Kx654CE | CYCLO HEXANE / CICLO ESANO , 0-100% LFL | 1294.38
| Kx654DE | R152a DIFLUOROETHANE detector , 0-100% LFL | 1294.38
| Kx654DT | DIMETHYL ETHER / DIMETILETERE , 0-100% LFL | 1294.38
| Kx654EN | ETHANE / ETANO , 0-100% LFL | 1294.38
| Kx654EP | HEPTANE / EPTANO , 0-100% LFL | 1294.38
| Kx654ES | HEXANE / ESANO , 0-100% LFL | 1294.38
| Kx654ET | ETHYL ALCOHOL / ALCOL ETILICO , 0-100% LFL | 1294.38
| Kx654GP | LPG / GPL , 0-100% LFL | 1294.38
| Kx654IB | ISO BUTANE / ISO BUTANO , 0-100% LFL | 1294.38
| Kx654IL | ISOBUTYLENE / ISOBUTILENE , 0-100% LFL | 1294.38
| Kx654IP | ISO PROPYL ALCOHOL / ALCOL ISO PROPILICO , 0-100% LFL | 1294.38
| Kx654PE | PENTANE / PENTANO , 0-100% LFL | 1294.38
| Kx654PP | PROPYLENE / PROPILENE , 0-100% LFL | 1294.38
| Kx654PR | PROPANE 7 PROPANO , 0-100% LFL | 1294.38
| Kx654TO | TOLUENE , 0-100% LFL | 1294.38
| Kx708HS | HYDROGEN SULPHIDE / IDROGENO SOLFORATO (H2S), 0-100 ppm | 863.46

## Gas Cylinders
Auto+Hide+Values
Gas Cylinder
| Product Code | Description | PVP
| SIB62 | Gas cylinder Oxygen O2/N2 - 17% Vol - 12L | 98.82
| SIB58 | Gas cylinder Carbon Monoxide CO/Air - 500 ppm - 12L | 98.82
| SIB571 | Gas cylinder Carbon Monoxide CO/Air - 300 ppm - 12L | 98.82
| SIB56 | Gas cylinder Carbon Monoxide CO/Air - 200 ppm - 12L | 98.82
| SIB512 | Gas cylinder R32/Air - 40% LFL - 12L | 98.82
| SIB508 | Gas cylinder R454C/Air - 30% LFL - 12L | 98.82
| SIB507 | Gas cylinder R454B/Air - 20% LFL - 12L | 98.82
| SIB505 | Gas cylinder R1234ZE/Air - 15% LFL - 12L | 98.82
| SIB504 | Gas cylinder R1234YF/Air - 15% LFL - 12L | 98.82
| SIB503 | Gas cylinder R32/Air - 30% LFL - 12L | 98.82
| SIB502 | Gas cylinder R1234ZE/Air - 30% LFL - 12L | 98.82
| SIB501 | Gas cylinder R1234YF/Air - 30% LFL - 12L | 98.82
| SIB50 | Gas cylinder Carbon Monoxide CO/Air - 55 ppm - 12L | 98.82
| SIB48 | Gas cylinder Carbon Dioxide CO2/N2 -25 % Vol - 12L | 98.82
| SIB44 | Gas cylinder Carbon Dioxide CO2/N2 -2 % Vol - 12L | 98.82
| SIB42 | Gas cylinder Carbon Dioxide CO2/N2 - 1 % Vol - 12L | 98.82
| SIB40 | Gas cylinder Carbon Dioxide CO2/N2 - 300 ppm - 12L | 98.82
| SIB36 | Gas cylinder Butane C4H10/Air - 30% LFL - 12L | 98.82
| SIB28 | Gas cylinder Propane C3H8/Air - 50% LFL - 12L | 98.82
| SIB26 | Gas cylinder Propane C3H8/Air - 34% LFL - 12L | 98.82
| SIB22 | Gas cylinder Propane C3H8/Air - 10% LFL - 12L | 98.82
| SIB19 | Gas cylinder Hydrogen H2/Air - 50% LFL - 12L | 98.82
| SIB09 | Gas cylinder Methane CH4/Air - 50% LFL - 12L | 98.82
| SIB06 | Gas cylinder Methane CH4/Air - 34% LFL - 12L | 98.82
| SIB021 | Gas cylinder Methane CH4/Air - 15% LFL - 12L | 98.82
| RIB93-34 | Gas cylinder SF6/Air - 1500 ppm - 34L | 427.68
| RIB92-34 | Gas cylinder H2/Air - 1000 ppm - 34L | 427.68
| RIB91-34 | Gas cylinder SO2/Air - 10 ppm - 34L | 427.68
| RIB82-34 | Gas cylinder Hydrogen Sulphide H2S/N2 - 10ppm - 34L | 427.68
| RIB81-34 | Gas cylinder Hydrogen Sulphide H2S/N2 - 40 ppm - 34L | 427.68
| RIB76-34 | Gas cylinder NH3/Air - 500 ppm - 34L | 427.68
| RIB74-34 | Gas cylinder NH3/Air - 100 ppm - 34L | 427.68
| RIB73-34 | Gas cylinder NH3/Air - 50 ppm - 34L | 427.68
| RIB70-34 | Gas cylinder NH3/Air - 5000 ppm - 34L | 427.68
| RIB151-34 | Gas cylinder R32/Air - 1500 ppm - 34L | 427.68
| RIB150-34 | Gas cylinder R507/Air - 1500 ppm - 34L | 427.68
| RIB149-34 | Gas cylinder R1234YF/N2 - 1500 ppm - 34L | 427.68
| RIB148-34 | Gas cylinder R134A/Air - 1500 ppm - 34L | 427.68
| RIB147-34 | Gas cylinder R410A/Air - 1500 ppm - 34L | 427.68
| RIB146-34 | Gas cylinder R1234ZE/Air - 1500 ppm - 34L | 427.68
| RIB145-34 | Gas cylinder SF6/Air - 1500 ppm - 34L | 427.68
| RIB144-34 | Gas cylinder R134/Air - 1500 ppm - 34L | 427.68
| RIB143-34 | Gas cylinder R407C/Air - 1500 ppm - 34L | 427.68
| RIB142-34 | Gas cylinder Isobutilene C4H8/Air - 1500 ppm - 34L | 427.68
| RIB141-34 | Gas cylinder Isobutylene C4H8/Air - 100 ppm - 34L | 427.68
| RIB140-34 | Gas cylinder Isobutylene C4H8/Air - 10 ppm - 34L | 427.68
| RIB138-34 | Gas cylinder R1234ZD/Air - 1500 ppm - 34L | 427.68
| RIB137-34 | Gas cylinder R448/Air - 1500 ppm - 34L | 427.68
| RIB136-34 | Gas cylinder R417A/Air - 1500 ppm - 34L | 427.68
| RIB135-34 | Gas cylinder R507A/Air - 1500 ppm - 34L | 427.68
| RIB134-34 | Gas cylinder R449/Air - 1500 ppm - 34L | 427.68
| RIB133-34 | Gas cylinder R407F/Air - 1500 ppm - 34L | 427.68
| RIB132-34 | Gas cylinder R407A/Air - 1500 ppm - 34L | 427.68
| RIB131-34 | Gas cylinder R404A/Air - 1500 ppm - 34L | 427.68
| RIB130-34 | Gas cylinder R125/Air - 1500 ppm - 34L | 427.68
| RIB129-34 | Gas cylinder SO2/Air - 50 ppm - 34L | 427.68
| RIB122-34 | Gas cylinder R410A/Air - 1000 ppm - 34L | 427.68
| RIB121-34 | Gas cylinder R1234YF/N2 - 1000 ppm - 34L | 427.68
| RIB117-34 | Gas cylinder Nitric Oxide NO/Air - 100 ppm - 34L | 427.68
| RIB116-34 | Gas cylinder Biossido di Azoto NO2/Air - 10 ppm - 34L | 427.68
| RIB110-34 | Gas cylinder Carbon Dioxide CO2/Air -2000 ppm - 34L | 427.68
| RIB109-34 | Gas cylinder R32/Air - 1500 ppm - 34L | 427.68
| RIB104-34 | Gas cylinder R1234ZE/Air - 1000 ppm - 34L | 427.68
| RIB103-34 | Gas cylinder R134A/Air - 1500 ppm - 34L | 427.68